## summary
ITEM | DESCRIPTION |  |  | MATLAB case | Remark | 2024-01-17 |  |  | 2024-01-18 |  |  | 2024-01-19
|  | Pass/Fail | Status % |  |  | MATLAB仿真状态 国标标准 | 矩泉测试状态 国标标准 | 锐能微测试状态 国标标准 | MATLAB仿真状态 国标标准 | 矩泉测试状态 国标标准 | 锐能微测试状态 国标标准 | MATLAB仿真状态 国标标准 | 矩泉测试状态 国标标准 | 锐能微测试状态 国标标准
UPD3688 System Evaluation
1 | 起动 |  |  | — | case不支持
2 | 潜动 |  |  | — | case不支持
3 | 基本误差 |  |  | case00 | case00覆盖，拆case(case00,基波电压幅度为0.4v（typ）) | Fail | Fail | Fail | Fail 全波有效值fail（PF=1,I=1/5000V，-1.4%；PF=0.5L,60Hz,I=1/5000V，-0.528%；PF=0.8C,I=1/5000V，-2.4%） 全波有功pass：-0.023% 全波无功pass：0.0084% 基波有效值fail（PF=1/0.5L/0.8C,I=1/5000V，1/2000V，-8.5926%） 基波有功fail（PF=1/0.5L/0.8C,I=1/5000V，1/2000V，1/1000V，-1.3569%） 基波无功pass：0.0541% | Fail 全波有功：0.12% | Fail 全波有功：0.19% | Fail 全波有效值fail（PF=1,I=1/5000V，-1.4%；PF=0.5L,60Hz,I=1/5000V，-0.528%；PF=0.8C,I=1/5000V，-2.4%） 全波有功pass：-0.023% 全波无功pass：0.0084% 基波有效值fail（PF=1/0.5L/0.8C,I=1/5000V，1/2000V，-8.5926%） 基波有功fail（PF=1/0.5L/0.8C,I=1/5000V，1/2000V，1/1000V，-1.3569%） 基波无功pass：0.0541% | Fail 全波有功：0.12% | Fail 全波有功：0.19%
4 | 计时准确度 |  |  | — | rtc，与emu无关
5 | 变差要求 |  |  | case03 | case03 |  | Pass | Pass | 未仿真结束 | Pass 全波有功误差改变量：0.0012% | Pass 全波有功误差改变量：0.0400% | Fail 全波有效值误差改变量fail（PF=1,I=1/80V，0.068%） 全波有功误差改变量fail（PF=1,I=1/80V，0.07%；PF=0.5L,I=1/80V，1/40V，0.136%） 全波无功误差改变量fail（PF=0.5L,I=1/80V，0.056%） 全波有效值误差改变量fail（PF=1/0,I=1/80V,0.068%） 全波有功误差改变量fail（基本全部fail，0.135%） 全波无功误差改变量fail（PF=0.5L/0,I=1/80V,0.06%） | Pass 全波有功误差改变量：0.0012% | Pass 全波有功误差改变量：0.0400%
6 | 负载电流升降变差 （此项测试电流上升、下降过程中在相同负载点的计量精度改变量，仿真做此测试无法呈现出与测试相同的效果） |  |  | case05 | 加case(从imin~imax：1->1/5000->1) |  | Pass | Fail |  | Pass 全波有功误差改变量：0.0072% | Fail 全波有功误差改变量：0.0345% |  | Pass 全波有功误差改变量：0.0072% | Fail 全波有功误差改变量：0.0345%
7 | 重复性 |  |  | — | 不用测
8 | 频率改变量-1Hz |  |  | case01 | case01(49hz~51hz/59hz~61hz)、case02(45hz~55hz/55hz~65hz) | Fail(case01) | Fail | Fail | Fail 全波有效值pass：-0.16% 全波有功fail（PF=1/0.5L，I=1/500V，-0.46%） 全波无功pass：-0.148% 基波有效值pass：-0.24% 基波有功fail（基本全部fail，-0.49%） 基波无功fail（FP=0.5L，频偏-1Hz，-0.2%） | Fail 全波有功：-0.15120% | Fail 全波有功：0.2% | Fail 全波有效值pass：-0.16% 全波有功fail（PF=1/0.5L，I=1/500V，-0.46%） 全波无功pass：-0.148% 基波有效值pass：-0.24% 基波有功fail（基本全部fail，-0.49%） 基波无功fail（FP=0.5L，频偏-1Hz，-0.2%） | Fail 全波有功：-0.15120% | Fail 全波有功：0.2%
| 频率改变量-2~5Hz |  |  | case02 | case01(49hz~51hz/59hz~61hz)、case02(45hz~55hz/55hz~65hz) |  | 未测试 | Fail | 未仿真结束 | 未测试 | Fail 全波有功：-0.2079% | Fail 全波有效值fail（PF=1/0.5L/0，I=1/2500V，-0.8%） 全波有功fail（PF=1，I=1/2500V，-0.9%；PF=0.5L，I=1/2500V，1/1000V，-1.6%） 全波无功fail（PF=0.5L/0，I=1/2500V，-1.3%） 基波有效值fail：-2.9% 基波有功fail：-1.7% 基波无功fail：-1.5% | 未测试 | Fail 全波有功：-0.2079%
9 | 电压改变量 |  |  | case04 | case00覆盖，拆case(case04，基波电压幅度为0.4v（0.6/0.8/1.2的倍数去扫）) |  | Fail | Fail | Fail 全波有效值pass：0.154% 全波有功pass:-0.058% 全波无功pass：-0.034% 基波有效值pass：-0.0491% 基波有功fail（PF=1/0.5L,I=1/500V,-0.1996%） 基波无功pass：-0.0381% | Fail 全波有功：-0.19300% | Fail 全波有功：+0.22450% | Fail 全波有效值pass：0.154% 全波有功pass:-0.058% 全波无功pass：-0.034% 基波有效值pass：-0.0491% 基波有功fail（PF=1/0.5L,I=1/500V,-0.1996%） 基波无功pass：-0.0381% | Fail 全波有功：-0.19300% | Fail 全波有功：+0.22450%
10 | 初始固有误差 |  |  | — | 不用测
11 | 负载电流快速改变 |  |  | case06 | 加case(开关交替) |  | Fail | Pass | 需要修改精度检查方式 此测试查看能量累加精度 | Fail 全波有功误差：-0.436% | Pass 全波有功误差：-0.025% | Fail 全波有效值：NA 全波有功fail：-0.7% 全波无功：NA 基波有效值：NA 基波有功fail：0.9% 基波无功：NA | Fail 全波有功误差：-0.436% | Pass 全波有功误差：-0.025%
12 | 电压、电流 中谐波-第五次谐波分量 |  |  | case16 | case16（只测5次），电压电流谐波同时加 | Pass 0.025% | Pass -0.04250% | Fail | Pass 全波有效值：0.0394% 全波有功：0.025% 全波无功：NA 基波有效值：0.0233% 基波有功：0.0525% 基波无功：NA | Pass 全波有功：-0.04250% | Fail 全波有功：-0.71810% | Pass 全波有效值：0.0394% 全波有功：0.025% 全波无功：NA 基波有效值：0.0233% 基波有功：0.0525% 基波无功：NA | Pass 全波有功：-0.04250% | Fail 全波有功：-0.71810%
13 | 电压、电流 中谐波-方顶波波形试验 |  |  | case10 | case10，电压电流谐波同时加 |  | Pass | Fail | Pass 全波有效值：0.0138% 全波有功：-0.0079% 全波无功：NA 基波有效值：0.0212% 基波有功：0.0532% 基波无功：NA | Pass 全波有功：+0.00470% | Fail 全波有功：+0.26830% | Pass 全波有效值：0.0138% 全波有功：-0.0079% 全波无功：NA 基波有效值：0.0212% 基波有功：0.0532% 基波无功：NA | Pass 全波有功：+0.00470% | Fail 全波有功：+0.26830%
14 | 电压、电流 中谐波-尖顶波波形试验 |  |  | case11 | case11，电压电流谐波同时加 |  | Pass | Fail | Pass 全波有效值：0.0268% 全波有功：-0.0095% 全波无功：NA 基波有效值：0.0211% 基波有功：0.542% 基波无功：NA | Pass 全波有功：+0.01720% | Fail 全波有功：+0.27540% | Pass 全波有效值：0.0268% 全波有功：-0.0095% 全波无功：NA 基波有效值：0.0211% 基波有功：0.542% 基波无功：NA | Pass 全波有功：+0.01720% | Fail 全波有功：+0.27540%
15 | 电流电路中间谐波-脉冲串触发波形试验 |  |  | case12 | case12，待讨论（上台体观测），只加电流谐波 |  | Pass | Pass | 测试激励已确认 需要修改精度检查方式 此测试查看能量累加精度 | Pass 全波有功：-0.00310% | Pass 全波有功：-0.03760% | Fail 全波有效值：NA 全波有功fail：-0.35% 全波无功：NA 基波有效值：NA 基波有功fail：1.1% 基波无功：NA | Pass 全波有功：-0.00310% | Pass 全波有功：-0.03760%
16 | 电流电路中的奇次谐波-90度相位触发波形试验 |  |  | case13 | case13，待讨论（上台体观测），只加电流谐波 |  | Pass | Pass | Pass 全波有效值：-0.2523% 全波有功：-0.0066% 全波无功：NA 基波有效值：-0.2644% 基波有功：0.0787% 基波无功：NA | Pass 全波有功：+0.01720% | Pass 全波有功：-0.02970% | Pass 全波有效值：-0.2523% 全波有功：-0.0066% 全波无功：NA 基波有效值：-0.2644% 基波有功：0.0787% 基波无功：NA | Pass 全波有功：+0.01720% | Pass 全波有功：-0.02970%
17 | 高次谐波试验 |  |  | case15 | case15（电压电流谐波分开加，15次~40次~15次扫） |  | 未测试完 | Pass | Fail 全波有效值fail（I=1/1000V，-0.56%，30~40次电压、电流谐波均fail） 全波有功pass：0.0747% 全波无功：NA 基波有效值pass：-0.27% 基波有功fail（I=1/1000V，-0.2791%，15~40次电压、电流谐波均fail） 基波无功：NA | 未测试完 | Pass 全波有功：-0.0173% | Fail 全波有效值fail（I=1/1000V，-0.56%，30~40次电压、电流谐波均fail） 全波有功pass：0.0747% 全波无功：NA 基波有效值pass：-0.27% 基波有功fail（I=1/1000V，-0.2791%，15~40次电压、电流谐波均fail） 基波无功：NA | 未测试完 | Pass 全波有功：-0.0173%
18 | 直流和偶次谐波-半波整流波形 |  |  | case14 | case14，只加电流谐波 |  | 没有相关组件，未测试 | 没有相关组件，未测试 | 测试激励已确认 需要修改精度检查方式 此测试查看能量累加精度 | 没有相关组件，未测试 | 没有相关组件，未测试 | Fail 全波有效值：NA 全波有功fail：-3.7% 全波无功：NA 基波有效值：NA 基波有功fail：4% 基波无功：NA | 没有相关组件，未测试 | 没有相关组件，未测试
19 | 环境温度变化 |  |  | — | 温度补偿-当前方案是模拟侧实现，需与模拟sync信息
20 | 功率测量精度 |  |  | — | 涵盖在case00中（基本误差测试涵盖）
21 | 电流有效值测量精度 |  |  | — | 涵盖在case00中（基本误差测试涵盖）
22 | 频率测量精度 |  |  | — | matlab case不支持过零点检测，需要外包继续提供有效matlab代码
23 | 功率因数测量精度 |  |  | — | matlab case不支持相位测量，需要外包继续提供有效matlab代码
24 | 温度测量精度 |  |  | — | rtc，与emu无关
25 | 全温下 RTC 计时精度 |  |  | — | rtc，与emu无关

## testcase-matlab
编号 | Description | 参数 | [14 empty] | 全波计量 |  |  |  | 基波计量
|  | ADC_samp_freq (ADC采样频率) | fund_wave_freq (基波频率) | ADC OSR | scan_freq (频率扫频范围) | PGA | fund_wave_amp_vol (基波电压幅度) | fund_wave_amp_current (基波电流幅度) | power_factor (功率因数) | vol_dc_bias （电压直流偏置） | current_dc_bias （电流直流偏置） | harmonic_wave_oder （输入谐波阶次） | amp_harmonic_vol (电压谐波幅度) | phase_harmonic_vol (电压谐波相位) | amp_harmonic_current (电流谐波幅度) | phase_harmonic_current (电流谐波相位) | full5_RMSU_err (RMSU电压有效值误差(‰)) | full5_RMSI_err (RMSI电流有效值误差(‰)) | full5_P_err (P有功功率误差(‰)) | full5_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰))
CASE00 | 基本误差 | 819200Hz（50Hz）/819200*1.2Hz（60Hz） | 50Hz/60Hz（typ） | 128/256 | NA | 1 | 0.4（typ） | 1/10000（摸底） 1/8000（摸底） 1/5000 1/2500 1/1000 1/400 1/200 1/20（typ） 1/10 1/2 1 | 1/0.5L/0.8C/0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 5 | 1 | 2 | 5 | 5 | 1 | 2
CASE01 | 频率改变量-1Hz | 819200Hz（50Hz）/819200*1.2Hz（60Hz） | 50Hz/60Hz（typ） | 128/256 | 49Hz:：2Hz：51Hz 59Hz:：2Hz：61Hz | 1 | 0.4（typ） | 1/5000（摸底） 1/2500 1/1000 1/500 1/20（typ） | 1/0.5L/0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 5 | 1 | 2 | 5 | 5 | 1 | 2
CASE02 | 频率改变量-5Hz | 819200Hz（50Hz）/819200*1.2Hz（60Hz） | 50Hz/60Hz（typ） | 128/256 | 45Hz:：3Hz：48Hz 52Hz:：3Hz：58Hz 62Hz:：3Hz：65Hz | 1 | 0.4（typ） | 1/5000（摸底） 1/2500 1/1000 1/500 1/20（typ） | 1/0.5L/0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 5 | 5 | 7 | 5 | 5 | 5 | 7
CASE03 | 变差要求 | 819200Hz（50Hz）/819200*1.2Hz（60Hz） | 50Hz/60Hz（typ） | 128/256 | NA | 1 | 0.4（typ） | 1/5000（摸底） 1/500（摸底） 1/80 1/40 | 1/0.5L/0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
CASE04 | 电压改变量 | 819200Hz（50Hz）/819200*1.2Hz（60Hz） | 50Hz/60Hz（typ） | 128/256 | NA | 1 | 0.4（typ）*（0.6（摸底）/0.8/1.2倍） | 1/5000（摸底） 1/2000（摸底） 1/1000 1/500 1/20（typ） 1 | 1/0.5L/0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 5 | 1 | 2 | 5 | 5 | 1 | 2
CASE05 | 负载电流升降变差 | 819200Hz（50Hz）/819200*1.2Hz（60Hz） | 50Hz/60Hz（typ） | 128/256 | NA | 1 | 0.4（typ） | 1/5000 1/2000 1/1000 1/400 1/200 1/20（typ） 1/10 1/2 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 5 | 1 |  | 5 | 5 | 1
CASE06 | 负载电流快速改变 看能量累加 | 819200Hz（50Hz）/819200*1.2Hz（60Hz） | 50Hz/60Hz（typ） | 128/256 | NA | 1 | 0.4（typ） | 1/40 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  | 1 |  |  |  | 1
CASE10 | 电压、电流 中谐波-方顶波波形试验 （电压电路同时加谐波） | 819200Hz（50Hz）/819200*1.2Hz（60Hz） | 50Hz/60Hz（typ） | 128/256 | NA | 1 | 0.4（typ） | 1/1000（摸底） 1/80（typ） 1/40 | 非1 | 0 | 0 | 3、5、7、11、13次 | 3.8%、2.4%、1.7%、1.0%、0.8% | 180°、180°、180°、180°、180° | 30%、18%、14%、9%、5% | 0、0、0、0、0 | 5 | 5 | 1 |  | 5 | 5 | 1
CASE11 | 电压、电流 中谐波-尖顶波波形试验 （电压电路同时加谐波） | 819200Hz（50Hz）/819200*1.2Hz（60Hz） | 50Hz/60Hz（typ） | 128/256 | NA | 1 | 0.4（typ） | 1/1000（摸底） 1/80（typ） 1/40 | 非1 | 0 | 0 | 3、5、7、11、13次 | 3.8%、2.4%、1.7%、1.0%、0.8% | 0、180°、0、0、180° | 30%、18%、14%、9%、5% | 180°、0、180°、180°、0 | 5 | 5 | 1 |  | 5 | 5 | 1
CASE12 | 谐波测试：脉冲串触发波形试验 （只加电流谐波，两个周波接通、两个周波关断） 看能量累加 | 819200Hz（50Hz）/819200*1.2Hz（60Hz） | 50Hz/60Hz（typ） | 128/256 | NA | 1 | 0.4（typ） | 1/1000（摸底） 1/80（typ） 1/40 | 1 | 0 | 0 | [7 empty] | 1 |  |  |  | 1
CASE13 | 谐波测试：90度相位触发波形试验 （只加电流谐波） | 819200Hz（50Hz）/819200*1.2Hz（60Hz） | 50Hz/60Hz（typ） | 128/256 | NA | 1 | 0.4（typ） | 1/1000（摸底） 1/80（typ） 1/40 | 非1 | 0 | 0 | 3：2：29 | 0 | 0 | 52%、18%、18%、10%、10%、7%、7%、5%、5%、3% | 0 | 5 | 5 | 1 |  | 5 | 5 | 1
CASE14 | 谐波测试：半波整流波形试验 （只加电流谐波） 看能量累加 | 819200Hz（50Hz）/819200*1.2Hz（60Hz） | 50Hz/60Hz（typ） | 128/256 | NA | 1 | 0.4（typ） | 1/8 1/4 | 1 | 0 | 0 | 2：2：20 | 0 | 0 | 42%、8%、4%、2%、1%、0.5%、0.25%、0.125%、0.0625%、0.03125% |  |  |  | 1 |  |  |  | 1
CASE15 | 谐波测试：高次谐波测试 （15~40次谐波） （电压电流谐波分开加） | 819200Hz（50Hz）/819200*1.2Hz（60Hz） | 50Hz/60Hz（typ） | 128/256 | NA | 1 | 0.4（typ） | 1/2000（摸底） 1/1000 1/20（typ） | 1 | 0 | 0 | 15：1：40 （每次只添加单次谐波） | 基波幅度2%(15~40次谐波） | 0 | 基波幅度10%(15~40次谐波） | 0 | 5 | 5 | 1 |  | 5 | 5 | 1
CASE16 | 电压、电流 中谐波-第五次谐波分量 （电压电路同时加谐波） | 819200Hz（50Hz）/819200*1.2Hz（60Hz） | 50Hz/60Hz（typ） | 128/256 | NA | 1 | 0.4（typ） | 1/1000（摸底） 1/20（typ） 1/10 | 1 | 0 | 0 | 5次 | 基波幅度10% | 0 | 基波幅度40% | 0 | 5 | 5 | 1 |  | 5 | 5 | 1

## 0117仿真结果VS锐能微校表台测试结果
误差单位% PGA增益16倍 锰铜采样电阻约250uΩ
Imin | Imax | Itr | Ib
0.25A | 60A | 0.5A | 5A
基本误差
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | 1 | 2 | 平均值 | 化整值 | 标准偏差 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
3 | 正向有功,基本误差 | 1.0 Imin | -0.1,0.1 | 2 | 1 | +0.0254 | +0.0173 | 0.0214 | 0.0 | 0.006 | 合格 |  | 1/1000
3 | 正向有功,基本误差 | 0.5L Imin | -0.1,0.1 | 2 | 1 | +0.1446 | +0.0872 | +0.11590 | +0.2 | 0.041 | 不合格 |  | 1/1000
3 | 正向有功,基本误差 | 0.8C Imin | -0.1,0.1 | 2 | 1 | -0.0523 | -0.0693 | -0.06080 | 0.0 | 0.012 | 合格 |  | 1/1000
3 | 正向有功,基本误差 | 1.0 Itr | -0.1,0.1 | 2 | 1 | +0.0230 | +0.0133 | +0.01820 | 0.0 | 0.007 | 合格 |  | 1/500
3 | 正向有功,基本误差 | 0.5L Itr | -0.1,0.1 | 2 | 1 | +0.1733 | +0.1590 | +0.16620 | +0.2 | 0.010 | 不合格 |  | 1/500
3 | 正向有功,基本误差 | 0.8C Itr | -0.1,0.1 | 2 | 1 | -0.0280 | -0.0377 | -0.03290 | 0.0 | 0.007 | 合格 |  | 1/500
3 | 正向有功,基本误差 | 1.0 10Itr | -0.1,0.1 | 2 | 3 | +0.0142 | +0.0150 | +0.01460 | 0.0 | 0.001 | 合格 |  | 1/50
3 | 正向有功,基本误差 | 0.5L 10Itr | -0.1,0.1 | 2 | 2 | +0.1715 | +0.1778 | +0.17470 | +0.2 | 0.004 | 不合格 |  | 1/50
3 | 正向有功,基本误差 | 0.8C 10Itr | -0.1,0.1 | 2 | 2 | -0.0525 | -0.0500 | -0.05130 | 0.0 | 0.002 | 合格 |  | 1/50
3 | 正向有功,基本误差 | 1.0 0.004Ib | -0.1,0.1 | 2 | 1 | +0.0736 | -0.0190 | +0.02730 | 0.0 | 0.065 | 合格 |  | 1/12500
3 | 正向有功,基本误差 | 1.0 0.01Ib | -0.1,0.1 | 2 | 1 | +0.0112 | +0.0361 | +0.02370 | 0.0 | 0.018 | 合格 |  | 1/5000
3 | 正向有功,基本误差 | 1.0 0.02Ib | -0.1,0.1 | 2 | 1 | +0.0708 | +0.0535 | +0.06220 | 0.0 | 0.012 | 合格 |  | 1/2500
3 | 正向有功,基本误差 | 0.5L 0.02Ib | -0.1,0.1 | 2 | 1 | +0.1700 | +0.1602 | +0.16510 | +0.2 | 0.007 | 不合格 |  | 1/2500
3 | 正向有功,基本误差 | 0.8C 0.02Ib | -0.1,0.1 | 2 | 1 | -0.0714 | -0.0747 | -0.07310 | 0.0 | 0.002 | 合格 |  | 1/2500
3 | 正向有功,基本误差 | 1.0 0.5Imax | -0.1,0.1 | 2 | 15 | +0.0250 | +0.0233 | +0.02420 | 0.0 | 0.001 | 合格 |  | 0.125
3 | 正向有功,基本误差 | 0.5L 0.5Imax | -0.1,0.1 | 2 | 13 | +0.1888 | +0.1927 | +0.19080 | +0.2 | 0.003 | 不合格 |  | 0.125
3 | 正向有功,基本误差 | 0.8C 0.5Imax | -0.1,0.1 | 2 | 13 | -0.0442 | -0.0442 | -0.04420 | 0.0 | 0.000 | 合格 |  | 0.125
3 | 正向有功,基本误差 | 1.0 Imax | -0.1,0.1 | 2 | 30 | +0.0267 | +0.0283 | +0.02750 | 0.0 | 0.001 | 合格 |  | 0.25
3 | 正向有功,基本误差 | 0.5L Imax | -0.1,0.1 | 2 | 25 | +0.1863 | +0.1843 | +0.18530 | +0.2 | 0.001 | 不合格 |  | 0.25
3 | 正向有功,基本误差 | 0.8C Imax | -0.1,0.1 | 2 | 25 | -0.0360 | -0.0360 | -0.03600 | 0.0 | 0.000 | 合格 |  | 0.25
变差要求
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | ErrorLoadTime | 1 | 2 | 平均值 | 化整值 | ErrorKeepTime | 误差改变量 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
3 | 正向有功,误差变差试验 | 1.0 Imin | -0.01,+0.01 | 2 | 1 | 00:01:00 | +0.0142 | -0.0044 | +0.00490 | 0.0 | 00:05:00 | 0.0186 | 不合格 |  | 1/1000
3 | 正向有功,误差变差试验 | 0.5L Imin | -0.01,+0.01 | 2 | 1 | 00:01:00 | +0.1235 | +0.1119 | +0.11770 | +0.2 | 00:05:00 | 0.0116 | 不合格 |  | 1/1000
3 | 正向有功,误差变差试验 | 1.0 Itr | -0.01,+0.01 | 2 | 1 | 00:01:00 | +0.0113 | +0.0022 | +0.00680 | 0.0 | 00:05:00 | 0.0091 | 合格 |  | 1/500
3 | 正向有功,误差变差试验 | 0.5L Itr | -0.01,+0.01 | 2 | 1 | 00:01:00 | +0.1459 | +0.1059 | +0.12590 | +0.2 | 00:05:00 | 0.04 | 不合格 |  | 1/500
3 | 正向有功,误差变差试验 | 1.0 10Itr | -0.01,+0.01 | 2 | 3 | 00:01:00 | +0.0117 | +0.0046 | +0.00820 | 0.0 | 00:05:00 | 0.0071 | 合格 |  | 1/50
3 | 正向有功,误差变差试验 | 0.5L 10Itr | -0.01,+0.01 | 2 | 2 | 00:01:00 | +0.1453 | +0.1308 | +0.13800 | +0.2 | 00:05:00 | 0.0145 | 不合格 |  | 1/50
负载电流升降变差
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | ErrorLoadTime | 误差1(%)升 | 误差2(%)升 | 误差3(%)升 | 误差4(%)升 | 误差5(%)升 | 平均误差(%)升 | 误差1(%)降 | 误差2(%)降 | 误差3(%)降 | 误差4(%)降 | 误差5(%)降 | 平均误差(%)降 | 平均值 | 化整值 | ErrorKeepTime | 误差改变量 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
3 | 正向有功,负载电流升降变差试验 | 1.0 Imin | -0.02,+0.02 | 2 | 1 | 00:00:00 | +0.0091 | -0.0096 | +0.0494 | +0.0005 | -0.0305 | +0.0038 | +0.0236 | +0.0486 | +0.0193 | +0.0480 | +0.0048 | +0.0289 | +0.01640 | 0.0 | 00:02:00 | 0.0251 | 不合格 |  | 1/1000
3 | 正向有功,负载电流升降变差试验 | 1.0 Itr | -0.02,+0.02 | 2 | 1 | 00:00:00 | -0.0002 | +0.0385 | +0.0193 | -0.0105 | +0.0253 | +0.0145 | +0.0308 | +0.0180 | +0.0058 | +0.0233 | +0.0288 | +0.0213 | +0.01790 | 0.0 | 00:02:00 | 0.0068 | 合格 |  | 1/500
3 | 正向有功,负载电流升降变差试验 | 1.0 10Itr | -0.02,+0.02 | 2 | 3 | 00:00:00 | +0.0225 | +0.0175 | +0.0225 | +0.0242 | +0.0142 | +0.0202 | +0.0242 | +0.0258 | +0.0283 | +0.0200 | +0.0233 | +0.0243 | +0.02220 | 0.0 | 00:02:00 | 0.0041 | 合格 |  | 1/50
3 | 正向有功,负载电流升降变差试验 | 1.0 0.01Ib | -0.02,+0.02 | 2 | 1 | 00:00:00 | +0.0669 | +0.0718 | -0.0363 | -0.0211 | +0.0498 | +0.0262 | +0.0289 | +0.0352 | -0.0519 | -0.0263 | -0.0275 | -0.0083 | +0.00900 | 0.0 | 00:02:00 | 0.0345 | 不合格 |  | 1/5000
3 | 正向有功,负载电流升降变差试验 | 1.0 0.02Ib | -0.02,+0.02 | 2 | 1 | 00:00:00 | -0.0254 | -0.0027 | +0.0677 | +0.0259 | +0.0187 | +0.0168 | -0.0422 | +0.0252 | +0.0139 | -0.0001 | +0.0240 | +0.0042 | +0.01050 | 0.0 | 00:02:00 | 0.0126 | 合格 |  | 1/2500
3 | 正向有功,负载电流升降变差试验 | 1.0 Imax | -0.02,+0.02 | 2 | 30 | 00:00:00 | +0.0333 | +0.0333 | +0.0350 | +0.0350 | +0.0367 | +0.0347 | +0.0250 | +0.0250 | +0.0250 | +0.0250 | +0.0233 | +0.0247 | +0.02970 | 0.0 | 00:02:00 | 0.0100 | 合格 |  | 0.25
频率改变量
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | 1 | 2 | 3 | 4 | 5 | 平均值 | 化整值 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
3 | 正向有功,抗其他频率改变试验 | 98%fn,1.0 Imin | -0.2,0.2 | 5 | 1 | +0.0507 | +0.0306 | +0.0423 | +0.0218 | +0.0384 | +0.03680 | 0.0 | 合格 |  | 1/1000
3 | 正向有功,抗其他频率改变试验 | 102%fn,1.0 Imin | -0.2,0.2 | 5 | 1 | -0.0636 | -0.0322 | -0.0124 | -0.0492 | -0.0405 | -0.03960 | 0.0 | 合格 |  | 1/1000
3 | 正向有功,抗其他频率改变试验 | 98%fn,0.5L Itr | -0.2,0.2 | 5 | 1 | +0.2421 | +0.2054 | +0.1866 | +0.1638 | +0.2225 | +0.20410 | +0.2 | 不合格 |  | 1/500
3 | 正向有功,抗其他频率改变试验 | 102%fn,0.5L Itr | -0.2,0.2 | 5 | 1 | +0.1650 | +0.1598 | +0.1334 | +0.1480 | +0.1477 | +0.15080 | +0.2 | 合格 |  | 1/500
3 | 正向有功,抗其他频率改变试验 | 102%fn,1.0 10Itr | -0.2,0.2 | 5 | 3 | -0.0516 | -0.0491 | -0.0525 | -0.0516 | -0.0508 | -0.05110 | 0.0 | 合格 |  | 1/50
3 | 正向有功,抗其他频率改变试验 | 98%fn,1.0 10Itr | -0.2,0.2 | 5 | 3 | +0.0292 | +0.0258 | +0.0250 | +0.0250 | +0.0267 | +0.02630 | 0.0 | 合格 |  | 1/50
3 | 正向有功,抗其他频率改变试验 | 98%fn,0.5L 10Itr | -0.2,0.2 | 5 | 2 | +0.2029 | +0.2004 | +0.2067 | +0.2029 | +0.2054 | +0.20370 | +0.2 | 不合格 |  | 1/50
3 | 正向有功,抗其他频率改变试验 | 102%fn,0.5L 10Itr | -0.2,0.2 | 5 | 2 | +0.1327 | +0.1377 | +0.1414 | +0.1452 | +0.1414 | +0.13970 | +0.2 | 合格 |  | 1/50
3 | 正向有功,抗其他频率改变试验 | 90%fn,1.0 0.01Ib | -0.2,0.2 | 5 | 1 | +0.0736 | +0.1565 | +0.1537 | +0.1057 | +0.1353 | +0.12500 | +0.2 | 合格 |  | 1/5000
3 | 正向有功,抗其他频率改变试验 | 110%fn,1.0 0.01Ib | -0.2,0.2 | 5 | 1 | -0.0989 | -0.1403 | -0.2077 | -0.1928 | -0.1478 | -0.15750 | -0.2 | 不合格 |  | 1/5000
3 | 正向有功,抗其他频率改变试验 | 90%fn,1.0 0.02Ib | -0.2,0.2 | 5 | 1 | +0.1441 | +0.0747 | +0.1655 | +0.1614 | +0.1723 | +0.14360 | +0.2 | 合格 |  | 1/2500
3 | 正向有功,抗其他频率改变试验 | 110%fn,1.0 0.02Ib | -0.2,0.2 | 5 | 1 | -0.1808 | -0.2147 | -0.2064 | -0.2375 | -0.2001 | -0.2079 | -0.2 | 不合格 |  | 1/2500
3 | 正向有功,抗其他频率改变试验 | 98%fn,1.0 Imax | -0.2,0.2 | 5 | 30 | +0.0250 | +0.0283 | +0.0283 | +0.0317 | +0.0317 | +0.02900 | 0.0 | 合格 |  | 0.25
3 | 正向有功,抗其他频率改变试验 | 102%fn,1.0 Imax | -0.2,0.2 | 5 | 30 | -0.0400 | -0.0400 | -0.0367 | -0.0367 | -0.0367 | -0.03800 | 0.0 | 合格 |  | 0.25
3 | 正向有功,抗其他频率改变试验 | 98%fn,0.5L Imax | -0.2,0.2 | 5 | 25 | +0.2044 | +0.2044 | +0.2084 | +0.2044 | +0.2064 | +0.20560 | +0.2 | 不合格 |  | 0.25
3 | 正向有功,抗其他频率改变试验 | 102%fn,0.5L Imax | -0.2,0.2 | 5 | 25 | +0.1422 | +0.1442 | +0.1442 | +0.1442 | +0.1442 | +0.14380 | +0.2 | 合格 |  | 0.25
电压改变量
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | 1 | 2 | 3 | 4 | 5 | 平均值 | 化整值 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
3 | 正向有功,抗其他电压改变试验 | 90%Un,1.0 Imin | -0.2,0.2 | 5 | 1 | -0.0080 | -0.0112 | +0.0048 | +0.0166 | +0.0009 | +0.00060 | 0.0 | 合格 |  | 1/1000
3 | 正向有功,抗其他电压改变试验 | 90%Un,1.0 10Itr | -0.2,0.2 | 5 | 3 | -0.0125 | -0.0125 | -0.0133 | -0.0142 | -0.0100 | -0.01250 | 0.0 | 合格 |  | 1/50
3 | 正向有功,抗其他电压改变试验 | 90%Un,1.0 Imax | -0.2,0.2 | 5 | 27 | -0.0074 | -0.0056 | -0.0019 | -0.0037 | -0.0019 | -0.00410 | 0.0 | 合格 |  | 0.25
3 | 正向有功,抗其他电压改变试验 | 90%Un,0.5L Itr | -0.2,0.2 | 5 | 1 | +0.1680 | +0.1447 | +0.1670 | +0.1447 | +0.1550 | +0.15590 | +0.2 | 合格 |  | 1/500
3 | 正向有功,抗其他电压改变试验 | 90%Un,0.5L 10Itr | -0.2,0.2 | 5 | 2 | +0.1440 | +0.1402 | +0.1440 | +0.1389 | +0.1402 | +0.14150 | +0.2 | 合格 |  | 1/50
3 | 正向有功,抗其他电压改变试验 | 90%Un,0.5L Imax | -0.2,0.2 | 5 | 25 | +0.1522 | +0.1542 | +0.1562 | +0.1562 | +0.1583 | +0.15540 | +0.2 | 合格 |  | 0.25
3 | 正向有功,抗其他电压改变试验 | 110%Un,1.0 Imin | -0.2,0.2 | 5 | 1 | -0.0049 | -0.0107 | +0.0039 | -0.0092 | -0.0094 | -0.00610 | 0.0 | 合格 |  | 1/1000
3 | 正向有功,抗其他电压改变试验 | 110%Un,1.0 10Itr | -0.2,0.2 | 5 | 3 | +0.0083 | +0.0075 | +0.0075 | +0.0042 | +0.0008 | +0.00570 | 0.0 | 合格 |  | 1/50
3 | 正向有功,抗其他电压改变试验 | 110%Un,1.0 Imax | -0.2,0.2 | 5 | 33 | +0.0136 | +0.0167 | +0.0152 | +0.0152 | +0.0167 | +0.01550 | 0.0 | 合格 |  | 0.25
3 | 正向有功,抗其他电压改变试验 | 110%Un,0.5L Itr | -0.2,0.2 | 5 | 1 | +0.2082 | +0.2250 | +0.2212 | +0.2413 | +0.2227 | +0.22370 | +0.2 | 不合格 |  | 1/500
3 | 正向有功,抗其他电压改变试验 | 110%Un,0.5L 10Itr | -0.2,0.2 | 5 | 2 | +0.2255 | +0.2130 | +0.2205 | +0.2117 | +0.2180 | +0.21770 | +0.2 | 不合格 |  | 1/50
3 | 正向有功,抗其他电压改变试验 | 110%Un,0.5L Imax | -0.2,0.2 | 5 | 25 | +0.2225 | +0.2245 | +0.2265 | +0.2245 | +0.2245 | 0.2245 | +0.2 | 不合格 |  | 0.25
3 | 正向有功,抗其他电压改变试验 | 80%Un,1.0 Itr | -0.2,0.2 | 5 | 1 | -0.0215 | -0.0257 | -0.0027 | -0.0002 | -0.0040 | -0.01080 | 0.0 | 合格 |  | 1/500
3 | 正向有功,抗其他电压改变试验 | 80%Un,1.0 10Itr | -0.2,0.2 | 5 | 2 | -0.0150 | -0.0175 | -0.0162 | -0.0225 | -0.0162 | -0.01750 | 0.0 | 合格 |  | 1/50
3 | 正向有功,抗其他电压改变试验 | 80%Un,1.0 Imax | -0.2,0.2 | 5 | 25 | -0.0100 | -0.0080 | -0.0060 | -0.0080 | -0.0080 | -0.00800 | 0.0 | 合格 |  | 0.25
3 | 正向有功,抗其他电压改变试验 | 85%Un,1.0 Itr | -0.2,0.2 | 5 | 1 | -0.0085 | -0.0027 | +0.0138 | -0.0265 | -0.0300 | -0.01080 | 0.0 | 合格 |  | 1/500
3 | 正向有功,抗其他电压改变试验 | 85%Un,1.0 10Itr | -0.2,0.2 | 5 | 3 | -0.0133 | -0.0150 | -0.0183 | -0.0108 | -0.0167 | -0.01480 | 0.0 | 合格 |  | 1/50
3 | 正向有功,抗其他电压改变试验 | 85%Un,1.0 Imax | -0.2,0.2 | 5 | 25 | -0.0080 | -0.0080 | -0.0080 | -0.0060 | -0.0060 | -0.00720 | 0.0 | 合格 |  | 0.25
3 | 正向有功,抗其他电压改变试验 | 115%Un,1.0 Itr | -0.2,0.2 | 5 | 1 | +0.0025 | +0.0425 | -0.0045 | +0.0095 | +0.0043 | +0.01090 | 0.0 | 合格 |  | 1/500
3 | 正向有功,抗其他电压改变试验 | 115%Un,1.0 10Itr | -0.2,0.2 | 5 | 3 | +0.0092 | +0.0100 | +0.0108 | +0.0125 | +0.0133 | +0.01120 | 0.0 | 合格 |  | 1/50
3 | 正向有功,抗其他电压改变试验 | 115%Un,1.0 Imax | -0.2,0.2 | 5 | 34 | +0.0191 | +0.0235 | +0.0250 | +0.0250 | +0.0279 | +0.02410 | 0.0 | 合格 |  | 0.25
负载电流快速改变
表位号 | 测试分类 | 项目名称 | 起始电量 | 电表脉冲数 | 标准表电量 | 终止电量 | 脉冲电量 | 脉冲电量误差 | Pulse_kWhRoundingError | 起止电量 | 起止电量误差 | Dial_kWhRoundingError | 结论 | 仿真结果 | 对应的DSM输入信号幅度
3 | 正向有功,抗其他负载电流快速改变试验 | 1.0 10Itr,10s,10s | False | 007025 | 4.39124 | False | 4.39062 | -0.014 | 0.0 | 0.0000 | -100.0000 | -100.0 | 不合格 |  | 1/50
3 | 正向有功,抗其他负载电流快速改变试验 | 1.0 10Itr,5s,5s | False | 007025 | 4.39094 | False | 4.39062 | -0.007 | 0.0 | 0.0000 | -100.0000 | -100.0 | 不合格 |  | 1/50
3 | 正向有功,抗其他负载电流快速改变试验 | 1.0 10Itr,5s,0.5s | False | 007023 | 4.39046 | False | 4.38938 | -0.025 | 0.0 | 0.0000 | -100.0000 | -100.0 | 不合格 |  | 1/50
电压、电流 中谐波-第五次谐波分量
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | 1 | 2 | 3 | 4 | 5 | 平均值 | 化整值 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
3 | 正向有功,谐波影响试验,谐波-第5次谐波试验 | 1.0 0.5Imax | -0.2,0.2 | 5 | 15 | -0.7181 | -0.7165 | -0.7198 | -0.7181 | -0.7181 | -0.7181 | -0.8 | 不合格 |  | 0.125
电压、电流 中谐波-方顶波波形试验
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | 1 | 2 | 平均值 | 化整值 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
3 | 正向有功,谐波影响试验,谐波-方顶波波形试验 | 方顶波,1.0 10Itr | -0.2,0.2 | 2 | 3 | +0.2691 | +0.2674 | 0.2683 | +0.2 | 不合格 |  | 1/50
电压、电流 中谐波-尖顶波波形试验
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | 1 | 2 | 3 | 4 | 5 | 平均值 | 化整值 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
3 | 正向有功,谐波影响试验,谐波-尖顶波波形试验 | 1.0 10Itr | -0.2,0.2 | 5 | 3 | +0.2774 | +0.2808 | 0.2724 | +0.2707 | +0.2758 | 0.2754 | +0.2 | 不合格 |  | 1/50
电流电路中间谐波-脉冲串触发波形试验
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | 1 | 2 | 3 | 4 | 5 | 平均值 | 化整值 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
3 | 正向有功,谐波影响试验,间谐波-脉冲串触发波形试验 | 1.0 10Itr | -0.2,0.2 | 5 | 3 | -0.0383 | -0.0083 | -0.0233 | -0.0550 | -0.0633 | -0.0376 | 0.0 | 合格 |  | 1/50
电流电路中的奇次谐波-90度相位触发波形试验
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | 1 | 2 | 3 | 4 | 5 | 平均值 | 化整值 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
3 | 正向有功,谐波影响试验,奇次谐波-90度相位触发波形试验 | 1.0 10Itr | -0.2,0.2 | 5 | 3 | -0.0267 | -0.0350 | -0.0267 | -0.0333 | -0.0267 | -0.0297 | 0.0 | 合格 |  | 1/50
高次谐波试验
表位号 | 测试分类 | 项目名称 | 上下限 | 影响量误差偏移极限值 | 结论 | 基本误差 | 最大误差 | 最小误差 | 误差改变量
3 | 正向有功,抗其他高次谐波试验 | 1.0 Itr | -0.2,+0.2 | 0.02 | 不合格 |  | +0.0180 | -0.0364 | -0.0364
|  | 仿真结果 | 对应的DSM输入信号幅度
U_15_1 | -0.0074 |  | 1/500
U_16_1 | +0.0053 |  | 1/500
U_17_1 | +0.0180 |  | 1/500
U_18_1 | -0.0072 |  | 1/500
U_19_1 | -0.0115 |  | 1/500
U_20_1 | -0.0059 |  | 1/500
U_21_1 | -0.0057 |  | 1/500
U_22_1 | -0.0055 |  | 1/500
U_23_1 | -0.0005 |  | 1/500
U_24_1 | -0.0154 |  | 1/500
U_25_1 | -0.0173 |  | 1/500
U_26_1 | -0.0290 |  | 1/500
U_27_1 | -0.0185 |  | 1/500
U_28_1 | -0.0188 |  | 1/500
U_29_1 | -0.0265 |  | 1/500
U_30_1 | -0.0275 |  | 1/500
U_31_1 | -0.0158 |  | 1/500
U_32_1 | -0.0215 |  | 1/500
U_33_1 | -0.0084 |  | 1/500
U_34_1 | -0.0107 |  | 1/500
U_35_1 | -0.0152 |  | 1/500
U_36_1 | -0.0224 |  | 1/500
U_37_1 | -0.0136 |  | 1/500
U_38_1 | -0.0191 |  | 1/500
U_39_1 | -0.0275 |  | 1/500
U_40_1 | -0.0357 |  | 1/500
U_40_2 | -0.0245 |  | 1/500
U_39_2 | -0.0120 |  | 1/500
U_38_2 | -0.0274 |  | 1/500
U_37_2 | -0.0259 |  | 1/500
U_36_2 | -0.0167 |  | 1/500
U_35_2 | -0.0255 |  | 1/500
U_34_2 | -0.0364 |  | 1/500
U_33_2 | -0.0219 |  | 1/500
U_32_2 | -0.0277 |  | 1/500
U_31_2 | -0.0170 |  | 1/500
U_30_2 | -0.0340 |  | 1/500
U_29_2 | -0.0270 |  | 1/500
U_28_2 | -0.0227 |  | 1/500
U_27_2 | -0.0322 |  | 1/500
U_26_2 | -0.0123 |  | 1/500
U_25_2 | -0.0127 |  | 1/500
U_24_2 | -0.0229 |  | 1/500
U_23_2 | -0.0201 |  | 1/500
U_22_2 | -0.0130 |  | 1/500
U_21_2 | -0.0223 |  | 1/500
U_20_2 | -0.0266 |  | 1/500
U_19_2 | -0.0183 |  | 1/500
U_18_2 | -0.0212 |  | 1/500
U_17_2 | -0.0044 |  | 1/500
U_16_2 | -0.0166 |  | 1/500
U_15_2 | -0.0168 |  | 1/500
I_15_1 | -0.0297 |  | 1/500
I_16_1 | -0.0215 |  | 1/500
I_17_1 | -0.0345 |  | 1/500
I_18_1 | -0.0157 |  | 1/500
I_19_1 | -0.0243 |  | 1/500
I_20_1 | -0.0268 |  | 1/500
I_21_1 | -0.0318 |  | 1/500
I_22_1 | -0.0100 |  | 1/500
I_23_1 | -0.0068 |  | 1/500
I_24_1 | -0.0258 |  | 1/500
I_25_1 | -0.0168 |  | 1/500
I_26_1 | -0.0148 |  | 1/500
I_27_1 | -0.0213 |  | 1/500
I_28_1 | -0.0160 |  | 1/500
I_29_1 | -0.0175 |  | 1/500
I_30_1 | -0.0153 |  | 1/500
I_31_1 | -0.0040 |  | 1/500
I_32_1 | -0.0257 |  | 1/500
I_33_1 | -0.0020 |  | 1/500
I_34_1 | -0.0210 |  | 1/500
I_35_1 | -0.0222 |  | 1/500
I_36_1 | -0.0153 |  | 1/500
I_37_1 | -0.0123 |  | 1/500
I_38_1 | -0.0082 |  | 1/500
I_39_1 | -0.0147 |  | 1/500
I_40_1 | -0.0080 |  | 1/500
I_40_2 | -0.0142 |  | 1/500
I_39_2 | -0.0190 |  | 1/500
I_38_2 | -0.0143 |  | 1/500
I_37_2 | -0.0162 |  | 1/500
I_36_2 | -0.0117 |  | 1/500
I_35_2 | -0.0147 |  | 1/500
I_34_2 | -0.0017 |  | 1/500
I_33_2 | -0.0023 |  | 1/500
I_32_2 | -0.0058 |  | 1/500
I_31_2 | -0.0215 |  | 1/500
I_30_2 | -0.0103 |  | 1/500
I_29_2 | -0.0093 |  | 1/500
I_28_2 | -0.0100 |  | 1/500
I_27_2 | +0.0115 |  | 1/500
I_26_2 | -0.0140 |  | 1/500
I_25_2 | -0.0080 |  | 1/500
I_24_2 | +0.0033 |  | 1/500
I_23_2 | -0.0083 |  | 1/500
I_22_2 | -0.0208 |  | 1/500
I_21_2 | -0.0193 |  | 1/500
I_20_2 | -0.0108 |  | 1/500
I_19_2 | -0.0080 |  | 1/500
I_18_2 | -0.0237 |  | 1/500
I_17_2 | -0.0085 |  | 1/500
I_16_2 | -0.0087 |  | 1/500
I_15_2 | -0.0265 |  | 1/500
直流和偶次谐波-半波整流波形

## 0117仿真结果VS钜泉校表台测试结果
误差单位% 互感线圈比例2500：1 电阻4Ω PGA2倍
Imin | Imax | Itr | Ib
0.25A | 60A | 0.5A | 5A
基本误差
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 平均值 | 化整值 | 标准偏差 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
1 | 正向有功,基本误差 | Lb:1.0 Imin | -0.2,0.2 | 10 | 1 | +0.0358 | +0.0315 | +0.0010 | +0.0170 | +0.0170 | +0.0250 | +0.0211 | +0.0367 | +0.0312 | +0.0128 | +0.02290 | 0.0 | 0.011 | 合格 |  | 0.00128 | 781.25
1 | 正向有功,基本误差 | Lb:0.8C Imin | -0.2,0.2 | 10 | 1 | -0.0049 | -0.0137 | -0.0068 | -0.0043 | -0.0184 | +0.0070 | +0.0066 | -0.0026 | -0.0177 | -0.0134 | -0.00680 | 0.0 | 0.009 | 合格 |  | 0.00128 | 781.25
1 | 正向有功,基本误差 | Lb:0.5L Imin | -0.2,0.2 | 10 | 1 | +0.0989 | +0.1106 | +0.0833 | +0.1069 | +0.0909 | +0.1118 | +0.1017 | +0.0996 | +0.0783 | +0.0979 | +0.09800 | 0.0 | 0.011 | 合格 |  | 0.00128 | 781.25
1 | 正向有功,基本误差 | Lb:1.0 Itr | -0.2,0.2 | 10 | 1 | -0.0012 | +0.0061 | -0.0016 | -0.0031 | -0.0003 | +0.0006 | +0.0011 | -0.0003 | +0.0016 | -0.0030 | 0.00000 | 0.0 | 0.003 | 合格 |  | 0.00256 | 390.625
1 | 正向有功,基本误差 | Lb:0.8C Itr | -0.2,0.2 | 10 | 1 | +0.0011 | -0.0197 | -0.0236 | -0.0336 | -0.0219 | -0.0411 | -0.0133 | -0.0281 | -0.0355 | -0.0262 | -0.02420 | 0.0 | 0.012 | 合格 |  | 0.00256 | 390.625
1 | 正向有功,基本误差 | Lb:0.5L Itr | -0.2,0.2 | 10 | 1 | +0.0566 | +0.0776 | +0.0780 | +0.0733 | +0.0830 | +0.0547 | +0.0624 | +0.0624 | +0.0655 | +0.0815 | +0.06950 | 0.0 | 0.010 | 合格 |  | 0.00256 | 390.625
1 | 正向有功,基本误差 | Lb:1.0 10Itr | -0.2,0.2 | 10 | 2 | -0.0109 | -0.0078 | -0.0070 | -0.0055 | -0.0102 | -0.0086 | -0.0102 | -0.0109 | -0.0094 | -0.0102 | -0.00910 | 0.0 | 0.002 | 合格 |  | 0.0256 | 39.0625
1 | 正向有功,基本误差 | Lb:0.8C 10Itr | -0.2,0.2 | 10 | 2 | -0.0258 | -0.0297 | -0.0250 | -0.0242 | -0.0242 | -0.0242 | -0.0227 | -0.0266 | -0.0227 | -0.0234 | -0.02490 | 0.0 | 0.002 | 合格 |  | 0.0256 | 39.0625
1 | 正向有功,基本误差 | Lb:0.5L 10Itr | -0.2,0.2 | 10 | 1 | +0.0188 | +0.0172 | +0.0172 | +0.0219 | +0.0156 | +0.0188 | +0.0172 | +0.0156 | +0.0219 | +0.0203 | +0.01850 | 0.0 | 0.002 | 合格 |  | 0.0256 | 39.0625
1 | 正向有功,基本误差 | Lb:1.0 0.5Imax | -0.2,0.2 | 10 | 10 | -0.0250 | -0.0266 | -0.0234 | -0.0234 | -0.0234 | -0.0234 | -0.0250 | -0.0234 | -0.0250 | -0.0219 | -0.02410 | 0.0 | 0.001 | 合格 |  | 0.16 | 6.25
1 | 正向有功,基本误差 | Lb:0.5L 0.5Imax | -0.2,0.2 | 10 | 8 | -0.0586 | -0.0547 | -0.0566 | -0.0566 | -0.0566 | -0.0547 | -0.0566 | -0.0527 | -0.0547 | -0.0527 | -0.05550 | 0.0 | 0.002 | 合格 |  | 0.16 | 6.25
1 | 正向有功,基本误差 | Lb:0.8C 0.5Imax | -0.2,0.2 | 10 | 8 | -0.0176 | -0.0176 | -0.0156 | -0.0176 | -0.0176 | -0.0176 | -0.0156 | -0.0195 | -0.0176 | -0.0195 | -0.01760 | 0.0 | 0.001 | 合格 |  | 0.16 | 6.25
1 | 正向有功,基本误差 | Lb:1.0 Imax | -0.2,0.2 | 10 | 19 | -0.0378 | -0.0411 | -0.0378 | -0.0395 | -0.0395 | -0.0395 | -0.0395 | -0.0395 | -0.0395 | -0.0395 | -0.03930 | 0.0 | 0.001 | 合格 |  | 0.32 | 3.125
1 | 正向有功,基本误差 | Lb:0.8C Imax | -0.2,0.2 | 10 | 16 | -0.0195 | -0.0215 | -0.0234 | -0.0195 | -0.0215 | -0.0195 | -0.0195 | -0.0195 | -0.0215 | -0.0176 | -0.02030 | 0.0 | 0.002 | 合格 |  | 0.32 | 3.125
1 | 正向有功,基本误差 | Lb:0.5L Imax | -0.2,0.2 | 10 | 16 | -0.1209 | -0.1229 | -0.1229 | -0.1209 | -0.1229 | -0.1248 | -0.1268 | -0.1209 | -0.1248 | -0.1248 | -0.12330 | -0.2 | 0.002 | 合格 |  | 0.32 | 3.125
1 | 正向有功,基本误差 | Lb:1.0 0.02Ib | -0.2,0.2 | 10 | 1 | +0.0765 | +0.0325 | +0.0552 | +0.0795 | +0.0662 | +0.0886 | +0.0486 | +0.0396 | +0.0631 | +0.0578 | +0.06080 | 0.0 | 0.018 | 合格 |  | 0.000512 | 1953.125
1 | 正向有功,基本误差 | Lb:1.0 0.01Ib | -0.2,0.2 | 10 | 1 | +0.0228 | +0.0961 | +0.0720 | +0.0699 | +0.0512 | +0.0549 | +0.0739 | +0.0497 | +0.0826 | +0.0429 | +0.06160 | 0.0 | 0.021 | 合格 |  | 0.000256 | 3906.25
1 | 正向有功,基本误差 | Lb:0.8C 0.02Ib | -0.2,0.2 | 10 | 1 | -0.0112 | -0.0079 | -0.0101 | +0.0289 | -0.0134 | -0.0356 | -0.0017 | +0.0098 | +0.0140 | +0.0110 | -0.00160 | 0.0 | 0.018 | 合格 |  | 0.000512 | 1953.125
1 | 正向有功,基本误差 | Lb:0.5L 0.02Ib | -0.2,0.2 | 10 | 1 | +0.0979 | +0.0770 | +0.1051 | +0.1076 | +0.0997 | +0.0695 | +0.1127 | +0.1051 | +0.0846 | +0.1391 | +0.09980 | 0.0 | 0.020 | 合格 |  | 0.000512 | 1953.125
变差要求
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | ErrorLoadTime | 1 | 2 | 平均值 | 化整值 | ErrorKeepTime | 误差改变量 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
1 | 正向有功,误差变差试验 | Lb:1.0 10Itr | -0.01,+0.01 | 2 | 2 | 00:01:00 | -0.0273 | -0.0285 | -0.02790 | 0.0 | 00:05:00 | 0.0012 | 合格 |  | 0.0256
1 | 正向有功,误差变差试验 | Lb:0.5L 10Itr | -0.01,+0.01 | 2 | 1 | 00:01:00 | +0.0086 | +0.0078 | +0.00820 | 0.0 | 00:05:00 | 0.0008 | 合格 |  | 0.0256
负载电流升降变差
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | ErrorLoadTime | 误差1(%)升 | 误差2(%)升 | 平均误差(%)升 | 误差1(%)降 | 误差2(%)降 | 平均误差(%)降 | 平均值 | 化整值 | ErrorKeepTime | 误差改变量 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
1 | 正向有功,负载电流升降变差试验 | Lb:1.0 Imin | -0.02,+0.02 | 2 | 1 | 00:00:00 | -0.0041 | -0.0111 | -0.0076 | -0.0039 | -0.0257 | -0.0148 | -0.01120 | 0.0 | 00:02:00 | 0.0072 | 合格 |  | 0.00128
1 | 正向有功,负载电流升降变差试验 | Lb:1.0 Itr | -0.02,+0.02 | 2 | 1 | 00:00:00 | -0.0320 | -0.0383 | -0.0352 | -0.0247 | -0.0337 | -0.0292 | -0.03220 | 0.0 | 00:02:00 | 0.0060 | 合格 |  | 0.00256
1 | 正向有功,负载电流升降变差试验 | Lb:1.0 10Itr | -0.02,+0.02 | 2 | 2 | 00:00:00 | -0.0398 | -0.0398 | -0.0398 | -0.0351 | -0.0328 | -0.0340 | -0.03690 | 0.0 | 00:02:00 | 0.0058 | 合格 |  | 0.0256
1 | 正向有功,负载电流升降变差试验 | Lb:1.0 Imax | -0.02,+0.02 | 2 | 19 | 00:00:00 | -0.0707 | -0.0707 | -0.0707 | -0.0641 | -0.0674 | -0.0658 | -0.06820 | 0.0 | 00:02:00 | 0.0049 | 合格 |  | 0.32
频率改变量
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | 1 | 2 | 平均值 | 化整值 | 结论 | 影响量误差偏移极限值 | 基本误差 | 误差改变量 | 仿真结果 | 对应的DSM输入信号幅度
1 | 正向有功,抗其他频率改变试验 | 98%fn,Lb:1.0 Imin | -0.2,0.2 | 2 | 1 | -0.0382 | -0.0389 | -0.03860 | 0.0 | 不合格 | 0.02 | +0.0124 | -0.0510 |  | 0.00128 | 781.25
1 | 正向有功,抗其他频率改变试验 | 98%fn,Lb:1.0 10Itr | -0.2,0.2 | 2 | 2 | -0.0609 | -0.0640 | -0.06250 | 0.0 | 不合格 | 0.02 | -0.0026 | -0.0599 |  | 0.0256 | 39.0625
1 | 正向有功,抗其他频率改变试验 | 98%fn,Lb:1.0 Imax | -0.2,0.2 | 2 | 19 | -0.0904 | -0.0904 | -0.09040 | 0.0 | 不合格 | 0.02 | -0.0340 | -0.0564 |  | 0.32 | 3.125
1 | 正向有功,抗其他频率改变试验 | 98%fn,Lb:0.5L Itr | -0.2,0.2 | 2 | 1 | +0.0289 | +0.0286 | +0.02880 | 0.0 | 不合格 | 0.02 | +0.0657 | -0.0369 |  | 0.00256 | 390.625
1 | 正向有功,抗其他频率改变试验 | 98%fn,Lb:0.5L 10Itr | -0.2,0.2 | 2 | 1 | -0.0328 | -0.0390 | -0.03590 | 0.0 | 不合格 | 0.02 | +0.0355 | -0.0714 |  | 0.0256 | 39.0625
1 | 正向有功,抗其他频率改变试验 | 98%fn,Lb:0.5L Imax | -0.2,0.2 | 2 | 16 | -0.1696 | -0.1735 | -0.17160 | -0.2 | 不合格 | 0.02 | -0.1268 | -0.0448 |  | 0.32 | 3.125
1 | 正向有功,抗其他频率改变试验 | 102%fn,Lb:1.0 Imin | -0.2,0.2 | 2 | 1 | -0.0176 | -0.0070 | -0.01230 | 0.0 | 不合格 | 0.02 | +0.0124 | -0.0247 |  | 0.00128 | 781.25
1 | 正向有功,抗其他频率改变试验 | 102%fn,Lb:1.0 10Itr | -0.2,0.2 | 2 | 2 | -0.0523 | -0.0547 | -0.05350 | 0.0 | 不合格 | 0.02 | -0.0026 | -0.0509 |  | 0.0256 | 39.0625
1 | 正向有功,抗其他频率改变试验 | 102%fn,Lb:1.0 Imax | -0.2,0.2 | 2 | 19 | -0.0805 | -0.0805 | -0.08050 | 0.0 | 不合格 | 0.02 | -0.0340 | -0.0465 |  | 0.32 | 3.125
1 | 正向有功,抗其他频率改变试验 | 102%fn,Lb:0.5L Itr | -0.2,0.2 | 2 | 1 | +0.0086 | +0.0116 | +0.01010 | 0.0 | 不合格 | 0.02 | +0.0657 | -0.0556 |  | 0.00256 | 390.625
1 | 正向有功,抗其他频率改变试验 | 102%fn,Lb:0.5L 10Itr | -0.2,0.2 | 2 | 1 | -0.0203 | -0.0172 | -0.01880 | 0.0 | 不合格 | 0.02 | +0.0355 | -0.0543 |  | 0.0256 | 39.0625
1 | 正向有功,抗其他频率改变试验 | 102%fn,Lb:0.5L Imax | -0.2,0.2 | 2 | 16 | -0.1502 | -0.1521 | -0.1512 | -0.2 | 不合格 | 0.02 | -0.1268 | -0.0244 |  | 0.32 | 3.125
电压改变量
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | 1 | 2 | 平均值 | 化整值 | 结论 | 影响量误差偏移极限值 | 基本误差 | 误差改变量 | 仿真结果 | 对应的DSM输入信号幅度
1 | 正向有功,抗其他电压改变试验 | 90%Un,Lb:1.0 Imin | -0.2,0.2 | 2 | 1 | -0.0466 | -0.0483 | -0.04750 | 0.0 | 不合格 | 0.03 | +0.0124 | -0.0599 |  | 0.00128
1 | 正向有功,抗其他电压改变试验 | 90%Un,Lb:1.0 10Itr | -0.2,0.2 | 2 | 2 | -0.0703 | -0.0679 | -0.06910 | 0.0 | 不合格 | 0.03 | -0.0026 | -0.0665 |  | 0.0256
1 | 正向有功,抗其他电压改变试验 | 90%Un,Lb:1.0 Imax | -0.2,0.2 | 2 | 17 | -0.1010 | -0.0992 | -0.10010 | -0.2 | 不合格 | 0.03 | -0.0340 | -0.0661 |  | 0.32
1 | 正向有功,抗其他电压改变试验 | 90%Un,Lb:0.5L Itr | -0.2,0.2 | 2 | 1 | -0.0039 | +0.0158 | +0.00600 | 0.0 | 不合格 | 0.03 | +0.0657 | -0.0597 |  | 0.00256
1 | 正向有功,抗其他电压改变试验 | 90%Un,Lb:0.5L 10Itr | -0.2,0.2 | 2 | 1 | -0.0578 | -0.0531 | -0.05550 | 0.0 | 不合格 | 0.03 | +0.0355 | -0.0910 |  | 0.0256
1 | 正向有功,抗其他电压改变试验 | 90%Un,Lb:0.5L Imax | -0.2,0.2 | 2 | 16 | -0.1930 | -0.1930 | -0.193 | -0.2 | 不合格 | 0.03 | -0.1268 | -0.0662 |  | 0.32
1 | 正向有功,抗其他电压改变试验 | 110%Un,Lb:1.0 Imin | -0.2,0.2 | 2 | 1 | -0.0365 | -0.0034 | -0.02000 | 0.0 | 不合格 | 0.03 | +0.0124 | -0.0324 |  | 0.00128
1 | 正向有功,抗其他电压改变试验 | 110%Un,Lb:1.0 10Itr | -0.2,0.2 | 2 | 2 | -0.0461 | -0.0437 | -0.04490 | 0.0 | 不合格 | 0.03 | -0.0026 | -0.0423 |  | 0.0256
1 | 正向有功,抗其他电压改变试验 | 110%Un,Lb:1.0 Imax | -0.2,0.2 | 2 | 21 | -0.0714 | -0.0729 | -0.07220 | 0.0 | 不合格 | 0.03 | -0.0340 | -0.0382 |  | 0.32
1 | 正向有功,抗其他电压改变试验 | 110%Un,Lb:0.5L Itr | -0.2,0.2 | 2 | 1 | +0.0025 | +0.0530 | +0.02780 | 0.0 | 不合格 | 0.03 | +0.0657 | -0.0379 |  | 0.00256
1 | 正向有功,抗其他电压改变试验 | 110%Un,Lb:0.5L 10Itr | -0.2,0.2 | 2 | 1 | -0.0187 | -0.0234 | -0.02110 | 0.0 | 不合格 | 0.03 | +0.0355 | -0.0566 |  | 0.0256
1 | 正向有功,抗其他电压改变试验 | 110%Un,Lb:0.5L Imax | -0.2,0.2 | 2 | 16 | -0.1541 | -0.1560 | -0.1551 | -0.2 | 合格 | 0.03 | -0.1268 | -0.0283 |  | 0.32
1 | 正向有功,抗其他电压改变试验 | 80%Un,Lb:1.0 Itr | -0.2,0.2 | 2 | 1 | -0.0726 | -0.0837 | -0.07820 | 0.0 | 不合格 | 0.03 | +0.0081 | -0.0863 |  | 0.00256
1 | 正向有功,抗其他电压改变试验 | 80%Un,Lb:1.0 10Itr | -0.2,0.2 | 2 | 2 | -0.0781 | -0.0781 | -0.07810 | 0.0 | 不合格 | 0.03 | -0.0026 | -0.0755 |  | 0.0256
1 | 正向有功,抗其他电压改变试验 | 80%Un,Lb:1.0 Imax | -0.2,0.2 | 2 | 16 | -0.1093 | -0.1093 | -0.10930 | -0.2 | 不合格 | 0.03 | -0.0340 | -0.0753 |  | 0.32
1 | 正向有功,抗其他电压改变试验 | 85%Un,Lb:1.0 Itr | -0.2,0.2 | 2 | 1 | -0.0748 | -0.0668 | -0.07080 | 0.0 | 不合格 | 0.03 | +0.0081 | -0.0789 |  | 0.00256
1 | 正向有功,抗其他电压改变试验 | 85%Un,Lb:1.0 10Itr | -0.2,0.2 | 2 | 2 | -0.0749 | -0.0726 | -0.07380 | 0.0 | 不合格 | 0.03 | -0.0026 | -0.0712 |  | 0.0256
1 | 正向有功,抗其他电压改变试验 | 85%Un,Lb:1.0 Imax | -0.2,0.2 | 2 | 16 | -0.1054 | -0.1054 | -0.10540 | -0.2 | 不合格 | 0.03 | -0.0340 | -0.0714 |  | 0.32
1 | 正向有功,抗其他电压改变试验 | 115%Un,Lb:1.0 Itr | -0.2,0.2 | 2 | 1 | -0.0269 | -0.0239 | -0.02540 | 0.0 | 不合格 | 0.03 | +0.0081 | -0.0335 |  | 0.00256
1 | 正向有功,抗其他电压改变试验 | 115%Un,Lb:1.0 10Itr | -0.2,0.2 | 2 | 2 | -0.0375 | -0.0375 | -0.03750 | 0.0 | 不合格 | 0.03 | -0.0026 | -0.0349 |  | 0.0256
1 | 正向有功,抗其他电压改变试验 | 115%Un,Lb:1.0 Imax | -0.2,0.2 | 2 | 22 | -0.0653 | -0.0653 | -0.06530 | 0.0 | 不合格 | 0.03 | -0.0340 | -0.0313 |  | 0.32
负载电流快速改变
表位号 | 测试分类 | 项目名称 | 起始电量 | 电表脉冲数 | 标准表电量 | 终止电量 | 脉冲电量 | 脉冲电量误差 | Pulse_kWhRoundingError | 起止电量 | 起止电量误差 | Dial_kWhRoundingError | 结论 | 仿真结果 | 对应的DSM输入信号幅度
1 | 正向有功,抗其他负载电流快速改变试验 | 1.0 10Itr,10s,10s | False | 004371 | 4.39015 | False | 4.37100 | -0.436 | -0.4 | 0.0000 | -100.0000 | -100.0 | 不合格 |  | 0.0256 | 39.0625
1 | 正向有功,抗其他负载电流快速改变试验 | 1.0 10Itr,5s,5s | False | 004372 | 4.39153 | False | 4.37200 | -0.445 | -0.4 | 0.0000 | -100.0000 | -100.0 | 不合格 |  | 0.0256
1 | 正向有功,抗其他负载电流快速改变试验 | 1.0 10Itr,5s,0.5s | False | 004371 | 4.39054 | False | 4.37100 | -0.445 | -0.4 | 0.0000 | -100.0000 | -100.0 | 不合格 |  | 0.0256
电压、电流 中谐波-第五次谐波分量
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | 1 | 2 | 平均值 | 化整值 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
1 | 正向有功,谐波影响试验,谐波-第5次谐波试验 | Lb:1.0 0.5Imax | -0.2,0.2 | 2 | 9 | -0.0416 | -0.0434 | -0.04250 | 0.0 | 合格 | 0.025 | 0.16
电压、电流 中谐波-方顶波波形试验
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | 1 | 2 | 平均值 | 化整值 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
1 | 正向有功,谐波影响试验,谐波-方顶波波形试验 | Lb:1.0 10Itr | -0.2,0.2 | 2 | 2 | +0.0047 | +0.0047 | 0.0047 | 0.0 | 合格 |  | 0.0256
电压、电流 中谐波-尖顶波波形试验
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | 1 | 2 | 平均值 | 化整值 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
1 | 正向有功,谐波影响试验,谐波-尖顶波波形试验 | Lb:1.0 10Itr | -0.2,0.2 | 2 | 2 | +0.0188 | +0.0156 | 0.0172 | 0.0 | 合格 |  | 0.0256
电流电路中间谐波-脉冲串触发波形试验
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | 1 | 平均值 | 化整值 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
1 | 正向有功,谐波影响试验,间谐波-脉冲串触发波形试验 | Lb:1.0 10Itr | -0.2,0.2 | 1 | 2 | -0.0031 | -0.0031 | 0.0 | 合格 |  |  | 0.0256
电流电路中的奇次谐波-90度相位触发波形试验
表位号 | 测试分类 | 项目名称 | 上下限 | 次数 | 误差脉冲圈数 | 1 | 2 | 平均值 | 化整值 | 结论 | 仿真结果 | 对应的DSM输入信号幅度
1 | 正向有功,谐波影响试验,奇次谐波-90度相位触发波形试验 | Lb:1.0 10Itr | -0.2,0.2 | 2 | 2 | +0.0172 | +0.0172 | 0.0172 | 0.0 | 合格 |  | 0.0256
高次谐波试验
表位号 | 测试分类 | 项目名称 | 上下限 | 影响量误差偏移极限值 | 结论 | 基本误差 | 最大误差 | 最小误差 | 误差改变量
1 | 正向有功,抗其他高次谐波试验 | Lb:1.0 Itr | -0.2,+0.2 | 0.02 | 不合格 | +0.0081 | -0.0257 | -0.0542 | -0.0623
|  | 仿真结果 | 对应的DSM输入信号幅度
U_15_1 | -0.0468 |  | 0.00256
U_16_1 | -0.0474 |  | 0.00256
U_17_1 | -0.0462 |  | 0.00256
U_18_1 | -0.0451 |  | 0.00256
U_19_1 | -0.0461 |  | 0.00256
U_20_1 | -0.0478 |  | 0.00256
U_21_1 | -0.0532 |  | 0.00256
U_22_1 | -0.0487 |  | 0.00256
U_23_1 | -0.0438 |  | 0.00256
U_24_1 | -0.0536 |  | 0.00256
U_25_1 | -0.0541 |  | 0.00256
U_26_1 | -0.0542 |  | 0.00256
U_27_1 | -0.0514 |  | 0.00256
U_28_1 | -0.0541 |  | 0.00256
U_29_1 | -0.0465 |  | 0.00256
U_30_1 | -0.0511 |  | 0.00256
U_31_1 | -0.0511 |  | 0.00256
U_32_1 | -0.0532 |  | 0.00256
U_33_1 | -0.0438 |  | 0.00256
U_34_1 | -0.0482 |  | 0.00256
U_35_1 | -0.0418 |  | 0.00256
U_36_1 | -0.0498 |  | 0.00256
U_37_1 | -0.0423 |  | 0.00256
U_38_1 | -0.0353 |  | 0.00256
U_39_1 | -0.0383 |  | 0.00256
U_40_1 | -0.0421 |  | 0.00256
U_40_2 | -0.0386 |  | 0.00256
U_39_2 | -0.0327 |  | 0.00256
U_38_2 | -0.0262 |  | 0.00256
U_37_2 | -0.0341 |  | 0.00256
U_36_2 | -0.0257 |  | 0.00256
直流和偶次谐波-半波整流波形

## osr=128校表系数
err取最后6组平均值
无补偿仿真结果 | 10：1 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.1 | 0.5 | 0 | 0 | -0.072 | -0.044 | -0.095 | -0.056
| 增益值 | [10 empty] | 1.0001 | 1 | 1.0001 | 1.0001
增益补偿仿真结果 | 5000：1 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | -0.028 | 81.613 | 22.067 | 12.487
| 偏置值 | [10 empty] | 3.55e-09 | 3.55e-09 | 4.4134e-07 | -4.3256e-07 | 3.55e-09 | 3.55e-09 | 4.4134e-07 | -4.3256e-07
无补偿仿真结果 | index | 参数 | [9 empty] | 全波计量 |  |  |  | 基波计量
|  | ADC_samp_freq (ADC采样频率) | fund_wave_freq (基波频率) | ADC OSR | scan_freq (频率扫频范围) | PGA | fund_wave_amp_vol (基波电压幅度) | fund_wave_amp_current (基波电流幅度) | power_factor (功率因数) | vol_dc_bias （电压直流偏置） | current_dc_bias （电流直流偏置） | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰))
| case0 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | -0.072 | -0.074 | -0.135 | -0.078 | -0.102 | -0.102 | -0.196 | -0.136
| case1 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.5 | 0.5 | 0 | 0 | -0.072 | -0.071 | -0.131 | -0.076 | -0.102 | -0.1 | -0.195 | -0.134
| case2 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.1 | 0.5 | 0 | 0 | -0.072 | -0.044 | -0.095 | -0.056 | -0.102 | -0.083 | -0.181 | -0.114
| case3 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | -0.072 | -0.009 | -0.051 | -0.031 | -0.102 | -0.062 | -0.164 | -0.088
| case4 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.005 | 0.5 | 0 | 0 | -0.072 | 0.731 | 0.747 | 0.433 | -0.102 | 0.328 | 0.147 | 0.365
| case5 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0025 | 0.5 | 0 | 0 | -0.072 | 1.783 | 1.639 | 0.924 | -0.102 | 0.768 | 0.498 | 0.868
| case6 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | -0.072 | 6.047 | 4.298 | 2.403 | -0.102 | 2.132 | 1.536 | 2.385
| case7 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 0.5 | 0 | 0 | -0.072 | 18.822 | 8.751 | 4.918 | -0.102 | 4.543 | 3.284 | 4.901
| case8 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | -0.072 | 81.541 | 21.967 | 12.413 | -0.102 | 12.833 | 8.402 | 12.52
增益补偿仿真结果 | index | 参数 | [9 empty] | 全波计量 |  |  |  | 基波计量
|  | ADC_samp_freq (ADC采样频率) | fund_wave_freq (基波频率) | ADC OSR | scan_freq (频率扫频范围) | PGA | fund_wave_amp_vol (基波电压幅度) | fund_wave_amp_current (基波电流幅度) | power_factor (功率因数) | vol_dc_bias （电压直流偏置） | current_dc_bias （电流直流偏置） | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰))
| case0 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | -0.028 | -0.03 | -0.04 | -0.022
| case1 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.5 | 0.5 | 0 | 0 | -0.028 | -0.027 | -0.036 | -0.02
| case2 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.1 | 0.5 | 0 | 0 | -0.028 | 0 | 0 | 0.001
| case3 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | -0.028 | 0.035 | 0.044 | 0.025
| case4 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.005 | 0.5 | 0 | 0 | -0.028 | 0.778 | 0.843 | 0.484
| case5 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0025 | 0.5 | 0 | 0 | -0.028 | 1.825 | 1.734 | 0.981
| case6 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | -0.028 | 6.093 | 4.394 | 2.452
| case7 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 0.5 | 0 | 0 | -0.028 | 18.854 | 8.848 | 4.973
| case8 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | -0.028 | 81.613 | 22.067 | 12.487
增益+偏置补偿仿真结果 | index | 参数 | [9 empty] | 全波计量 |  |  |  | 基波计量
|  | ADC_samp_freq (ADC采样频率) | fund_wave_freq (基波频率) | ADC OSR | scan_freq (频率扫频范围) | PGA | fund_wave_amp_vol (基波电压幅度) | fund_wave_amp_current (基波电流幅度) | power_factor (功率因数) | vol_dc_bias （电压直流偏置） | current_dc_bias （电流直流偏置） | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰))
| case0 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | -0.028 | -0.03 | -0.044 | -0.024 | -0.058 | -0.058 | -0.106 | -0.083
| case1 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.5 | 0.5 | 0 | 0 | -0.028 | -0.027 | -0.044 | -0.025 | -0.058 | -0.056 | -0.109 | -0.083
| case2 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.1 | 0.5 | 0 | 0 | -0.028 | 0 | -0.044 | -0.025 | -0.058 | -0.04 | -0.13 | -0.083
| case3 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | -0.028 | 0.033 | -0.044 | -0.025 | -0.058 | -0.019 | -0.157 | -0.082
| case4 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.005 | 0.5 | 0 | 0 | -0.028 | 0.635 | -0.04 | -0.01 | -0.058 | 0.229 | -0.64 | -0.078
| case5 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0025 | 0.5 | 0 | 0 | -0.028 | 1.256 | -0.031 | -0.019 | -0.058 | 0.244 | -1.172 | -0.074
| case6 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | -0.028 | 2.556 | -0.02 | -0.043 | -0.058 | -1.373 | -2.783 | -0.056
| case7 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 0.5 | 0 | 0 | -0.028 | 4.816 | 0.023 | -0.022 | -0.058 | -9.651 | -5.447 | -0.037
| case8 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | -0.028 | -4.089 | -0.001 | -0.01 | -0.058 | -78.964 | -13.57 | 0.091

## osr=256校表系数
err取最后6组平均值
无补偿仿真结果 | 10：1 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.1 | 0.5 | 0 | 0 | 0.046 | 0.041 | 0.068 | 0.078
| 增益值 | [10 empty] | 0.99995 | 0.99996 | 0.99993 | 0.99992
增益补偿仿真结果 | 5000：1 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 0.005 | 10.361 | -13.678 | -1.62
| 偏置值 | [10 empty] | 4.1662e-10 | 4.1662e-10 | -2.7356e-07 | 5.6118e-08 | 4.1662e-10 | 4.1662e-10 | -2.7356e-07 | 5.6118e-08
无补偿仿真结果 | index | 参数 | [9 empty] | 全波计量 |  |  |  | 基波计量
|  | ADC_samp_freq (ADC采样频率) | fund_wave_freq (基波频率) | ADC OSR | scan_freq (频率扫频范围) | PGA | fund_wave_amp_vol (基波电压幅度) | fund_wave_amp_current (基波电流幅度) | power_factor (功率因数) | vol_dc_bias （电压直流偏置） | current_dc_bias （电流直流偏置） | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰))
| case0 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | 0.046 | 0.049 | 0.092 | 0.081 | 0.294 | 0.296 | 0.587 | -0.089
| case1 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.5 | 0.5 | 0 | 0 | 0.046 | 0.048 | 0.09 | 0.081 | 0.294 | 0.295 | 0.584 | -0.089
| case2 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.1 | 0.5 | 0 | 0 | 0.046 | 0.041 | 0.068 | 0.078 | 0.294 | 0.288 | 0.567 | -0.091
| case3 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 0.046 | 0.033 | 0.041 | 0.075 | 0.294 | 0.281 | 0.544 | -0.094
| case4 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.005 | 0.5 | 0 | 0 | 0.046 | -0.1 | -0.447 | 0.013 | 0.294 | 0.141 | 0.147 | -0.137
| case5 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0025 | 0.5 | 0 | 0 | 0.046 | -0.193 | -0.985 | -0.055 | 0.294 | -0.012 | -0.292 | -0.186
| case6 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 0.046 | -0.213 | -2.619 | -0.255 | 0.294 | -0.458 | -1.627 | -0.327
| case7 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 0.5 | 0 | 0 | 0.046 | 0.715 | -5.327 | -0.591 | 0.294 | -1.158 | -3.838 | -0.566
| case8 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 0.046 | 10.397 | -13.611 | -1.541 | 0.294 | -2.928 | -10.648 | -1.211
增益补偿仿真结果 | index | 参数 | [9 empty] | 基波计量 |  |  |  | 基波计量
|  | ADC_samp_freq (ADC采样频率) | fund_wave_freq (基波频率) | ADC OSR | scan_freq (频率扫频范围) | PGA | fund_wave_amp_vol (基波电压幅度) | fund_wave_amp_current (基波电流幅度) | power_factor (功率因数) | vol_dc_bias （电压直流偏置） | current_dc_bias （电流直流偏置） | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰))
| case0 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | 0.005 | 0.008 | 0.024 | 0.003 | 0.006 | 0.008 | 0.02 | 0.002
| case1 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.5 | 0.5 | 0 | 0 | 0.005 | 0.007 | 0.022 | 0.003 | 0.006 | 0.007 | 0.017 | 0.002
| case2 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.1 | 0.5 | 0 | 0 | 0.005 | 0 | 0 | 0 | 0.006 | 0 | 0 | 0
| case3 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 0.005 | -0.008 | -0.027 | -0.003 | 0.006 | -0.007 | -0.023 | -0.003
| case4 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.005 | 0.5 | 0 | 0 | 0.005 | -0.141 | -0.514 | -0.065 | 0.006 | -0.147 | -0.42 | -0.046
| case5 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0025 | 0.5 | 0 | 0 | 0.005 | -0.235 | -1.053 | -0.133 | 0.006 | -0.3 | -0.859 | -0.095
| case6 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 0.005 | -0.25 | -2.687 | -0.333 | 0.006 | -0.747 | -2.193 | -0.236
| case7 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 0.5 | 0 | 0 | 0.005 | 0.675 | -5.395 | -0.669 | 0.006 | -1.446 | -4.403 | -0.475
| case8 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 0.005 | 10.361 | -13.678 | -1.62 | 0.006 | -3.215 | -11.208 | -1.121
增益+偏置补偿仿真结果 | index | 参数 | [9 empty] | 基波计量 |  |  |  | 基波计量
|  | ADC_samp_freq (ADC采样频率) | fund_wave_freq (基波频率) | ADC OSR | scan_freq (频率扫频范围) | PGA | fund_wave_amp_vol (基波电压幅度) | fund_wave_amp_current (基波电流幅度) | power_factor (功率因数) | vol_dc_bias （电压直流偏置） | current_dc_bias （电流直流偏置） | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰))
| case0 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | 0.005 | 0.008 | 0.027 | 0.004 | 0.253 | 0.255 | 0.521 | -0.166
| case1 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.5 | 0.5 | 0 | 0 | 0.005 | 0.007 | 0.027 | 0.004 | 0.253 | 0.254 | 0.522 | -0.166
| case2 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.1 | 0.5 | 0 | 0 | 0.005 | 0 | 0.027 | 0.003 | 0.253 | 0.247 | 0.526 | -0.166
| case3 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 0.005 | -0.009 | 0.027 | 0.003 | 0.253 | 0.239 | 0.531 | -0.165
| case4 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.005 | 0.5 | 0 | 0 | 0.005 | -0.157 | 0.033 | 0 | 0.253 | 0.083 | 0.626 | -0.151
| case5 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0025 | 0.5 | 0 | 0 | 0.005 | -0.301 | 0.041 | -0.003 | 0.253 | -0.12 | 0.734 | -0.134
| case6 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 0.005 | -0.667 | 0.049 | -0.009 | 0.253 | -0.917 | 1.04 | -0.081
| case7 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 0.5 | 0 | 0 | 0.005 | -0.99 | 0.077 | -0.021 | 0.253 | -2.869 | 1.565 | 0.004
| case8 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 0.005 | 0 | 0.001 | 0.001 | 0.253 | -13.471 | 2.964 | 0.331

## case00 test report
dynamic_test | ADC_samp_freq | fund_wave_freq | ADC_OSR | scan_freq | PGA | fund_wave_amp_vol | fund_wave_amp_current | power_factor | vol_dc_bias | current_dc_bias | harmonic_wave_oder | amp_harmonic_vol | amp_harmonic_current | full_RMSU_err (RMSU电压有效值误差(‰)) | full_RMSI_err (RMSI电流有效值误差(‰)) | full_P_err (P有功功率误差(‰)) | full_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | 备注
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.031 | -0.054 |  | -0.058 | -0.06 | -0.113
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.5 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.029 | -0.054 |  | -0.058 | -0.059 | -0.115
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.012 | -0.054 |  | -0.058 | -0.052 | -0.125
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.01 | -0.054 |  | -0.058 | -0.045 | -0.139
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.005 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.412 | -0.052 |  | -0.058 | -0.025 | -0.38
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0025 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.809 | -0.049 |  | -0.058 | -0.267 | -0.647
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 1.51 | -0.041 |  | -0.058 | -2.656 | -1.452
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 2.15 | -0.025 |  | -0.058 | -12.235 | -2.789
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -13.676 | 0.028 |  | -0.058 | -85.916 | -6.792
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.000125 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -44.61 | 0.077 |  | -0.058 | -244.93 | -10.787 |  | 摸底
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -89.233 | 0.114 |  | -0.058 | -433.364 | -13.464 |  | 摸底
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.008 | 0.028 |  | 0.253 | 0.254 | 0.523
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.5 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.006 | 0.028 |  | 0.253 | 0.253 | 0.523
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.005 | 0.028 |  | 0.253 | 0.244 | 0.525
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.005 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.256 | 0.03 |  | 0.253 | 0.018 | 0.574
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0025 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.498 | 0.033 |  | 0.253 | -0.25 | 0.627
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -1.16 | 0.042 |  | 0.253 | -1.241 | 0.784
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -1.998 | 0.056 |  | 0.253 | -3.522 | 1.047
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -2.233 | 0.1 |  | 0.253 | -15.137 | 1.839
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.000125 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.275 | 0.154 |  | 0.253 | -34.129 | 2.639 |  | 摸底
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0001 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | 2.4 | 0.196 |  | 0.253 | -51.11 | 3.177 |  | 摸底
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.031 | -0.054 |  | -0.059 | -0.06 | -0.115
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.5 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.029 | -0.054 |  | -0.059 | -0.06 | -0.116
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.011 | -0.054 |  | -0.059 | -0.053 | -0.127
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.011 | -0.054 |  | -0.059 | -0.046 | -0.141
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.005 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.406 | -0.051 |  | -0.059 | -0.026 | -0.382
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0025 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.788 | -0.049 |  | -0.059 | -0.268 | -0.649
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 1.458 | -0.04 |  | -0.059 | -2.655 | -1.452
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 2.047 | -0.026 |  | -0.059 | -12.236 | -2.791
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -8.801 | 0.022 |  | -0.059 | -85.926 | -6.802
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.000125 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -31.535 | 0.084 |  | -0.059 | -244.922 | -10.782 |  | 摸底
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -84.372 | 0.115 |  | -0.059 | -433.372 | -13.469 |  | 摸底
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.008 | 0.029 |  | 0.253 | 0.254 | 0.522
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.5 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.006 | 0.028 |  | 0.253 | 0.253 | 0.522
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.005 | 0.028 |  | 0.253 | 0.244 | 0.524
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.005 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.255 | 0.03 |  | 0.253 | 0.018 | 0.574
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0025 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.5 | 0.033 |  | 0.253 | -0.25 | 0.627
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -1.161 | 0.041 |  | 0.253 | -1.242 | 0.784
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -1.988 | 0.054 |  | 0.253 | -3.525 | 1.044
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -2.285 | 0.095 |  | 0.253 | -15.142 | 1.833
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.000125 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.528 | 0.149 |  | 0.253 | -34.133 | 2.635 |  | 摸底
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0001 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | 1.975 | 0.198 |  | 0.253 | -51.108 | 3.178 |  | 摸底
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.03 | -0.044 | -0.024 | -0.058 | -0.058 | -0.106 | -0.083
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.5 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.027 | -0.044 | -0.025 | -0.058 | -0.056 | -0.109 | -0.083
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.1 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0 | -0.044 | -0.025 | -0.058 | -0.04 | -0.13 | -0.083
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.033 | -0.044 | -0.024 | -0.058 | -0.019 | -0.157 | -0.082
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.005 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.634 | -0.04 | -0.014 | -0.058 | 0.23 | -0.64 | -0.078
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0025 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | 1.255 | -0.031 | -0.019 | -0.058 | 0.244 | -1.172 | -0.075
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | 2.552 | -0.02 | -0.045 | -0.058 | -1.373 | -2.782 | -0.056
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | 4.804 | 0.022 | -0.02 | -0.058 | -9.652 | -5.448 | -0.037
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | -4.067 | -0.003 | -0.023 | -0.058 | -78.964 | -13.569 | 0.09
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.000125 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | -34.661 | -0.128 | 0.077 | -0.058 | -231.569 | -21.817 | 0.254 | 摸底
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0001 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | -85.941 | -0.079 | 0.025 | -0.058 | -410.743 | -27.179 | 0.304 | 摸底
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.008 | 0.027 | 0.004 | 0.253 | 0.255 | 0.521 | -0.166
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.5 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.007 | 0.027 | 0.004 | 0.253 | 0.254 | 0.522 | -0.166
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.1 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0 | 0.027 | 0.003 | 0.253 | 0.247 | 0.526 | -0.166
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.009 | 0.027 | 0.003 | 0.253 | 0.239 | 0.531 | -0.165
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.005 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.157 | 0.033 | 0 | 0.253 | 0.083 | 0.626 | -0.15
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0025 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.301 | 0.041 | -0.003 | 0.253 | -0.12 | 0.734 | -0.134
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.669 | 0.049 | -0.009 | 0.253 | -0.916 | 1.04 | -0.081
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.991 | 0.077 | -0.021 | 0.253 | -2.869 | 1.565 | 0.004
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.003 | 0 | 0 | 0.253 | -13.472 | 2.963 | 0.331
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.000125 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | 3.756 | -0.026 | 0.027 | 0.253 | -31.393 | 4.445 | 0.642 | 摸底
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0001 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | 7.574 | 0.01 | 0.005 | 0.253 | -47.638 | 5.476 | 0.817 | 摸底
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.03 | -0.042 | -0.024 | -0.059 | -0.059 | -0.108 | -0.084
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.5 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.026 | -0.043 | -0.025 | -0.059 | -0.057 | -0.111 | -0.084
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.1 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0 | -0.042 | -0.024 | -0.059 | -0.04 | -0.132 | -0.084
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.033 | -0.042 | -0.025 | -0.059 | -0.02 | -0.159 | -0.084
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.005 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.635 | -0.038 | -0.022 | -0.059 | 0.229 | -0.642 | -0.08
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0025 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | 1.268 | -0.03 | -0.018 | -0.059 | 0.243 | -1.175 | -0.076
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | 2.564 | -0.018 | -0.029 | -0.059 | -1.374 | -2.784 | -0.058
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | 4.502 | 0.018 | -0.084 | -0.059 | -9.653 | -5.45 | -0.039
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | -5.28 | 0.014 | -0.075 | -0.059 | -78.971 | -13.562 | 0.078
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.000125 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | -36.969 | -0.106 | 0.162 | -0.059 | -231.569 | -21.814 | 0.254 | 摸底
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0001 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.028 | -73.828 | -0.046 | 0.172 | -0.059 | -410.721 | -27.154 | 0.313 | 摸底
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.008 | 0.028 | 0.004 | 0.253 | 0.255 | 0.521 | -0.167
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.5 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.007 | 0.028 | 0.004 | 0.253 | 0.254 | 0.522 | -0.167
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.1 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0 | 0.028 | 0.003 | 0.253 | 0.247 | 0.526 | -0.166
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.008 | 0.028 | 0.003 | 0.253 | 0.239 | 0.531 | -0.166
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.005 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.157 | 0.033 | 0 | 0.253 | 0.083 | 0.626 | -0.151
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0025 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.299 | 0.041 | -0.003 | 0.253 | -0.12 | 0.733 | -0.135
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.668 | 0.05 | -0.01 | 0.253 | -0.917 | 1.04 | -0.082
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | -1.017 | 0.078 | -0.021 | 0.253 | -2.868 | 1.567 | 0.004
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.261 | 0 | 0.002 | 0.253 | -13.472 | 2.962 | 0.332
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.000125 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | 3.28 | -0.029 | 0.021 | 0.253 | -31.398 | 4.443 | 0.636 | 摸底
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0001 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | 7.425 | 0.002 | 0.003 | 0.253 | -47.643 | 5.465 | 0.814 | 摸底
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 1 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.033 | -0.059 | -0.014 | -0.058 | -0.06 | -0.114 | -0.072
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.5 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.033 | -0.059 | -0.014 | -0.058 | -0.06 | -0.112 | -0.072
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.1 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.035 | -0.059 | -0.015 | -0.058 | -0.056 | -0.099 | -0.072
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.038 | -0.059 | -0.015 | -0.058 | -0.052 | -0.082 | -0.072
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.005 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.065 | -0.064 | -0.031 | -0.058 | -0.102 | 0.218 | -0.066
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0025 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.172 | -0.069 | -0.011 | -0.058 | -0.42 | 0.551 | -0.057
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -1.032 | -0.077 | -0.051 | -0.058 | -3.04 | 1.557 | -0.04
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -4.093 | -0.101 | -0.01 | -0.058 | -13.027 | 3.224 | 0.008
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0002 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -23.598 | -0.227 | -0.028 | -0.058 | -88.252 | 8.166 | 0.197
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.000125 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -50.994 | -0.337 | 0.248 | -0.058 | -250.199 | 13.145 | 0.278 | 摸底
case00 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0001 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -101.642 | -0.383 | -0.214 | -0.058 | -445.191 | 16.517 | 0.286 | 摸底
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 1 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.01 | 0.029 | 0.002 | 0.253 | 0.256 | 0.523 | -0.167
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.5 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.011 | 0.029 | 0.002 | 0.253 | 0.257 | 0.522 | -0.167
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.1 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.018 | 0.029 | 0.002 | 0.253 | 0.263 | 0.519 | -0.166
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.028 | 0.028 | 0.002 | 0.253 | 0.27 | 0.516 | -0.165
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.005 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.207 | 0.024 | -0.002 | 0.253 | 0.384 | 0.456 | -0.143
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0025 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.422 | 0.02 | -0.004 | 0.253 | 0.482 | 0.389 | -0.117
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 1.162 | 0.01 | -0.02 | 0.253 | 0.588 | 0.193 | -0.048
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 2.619 | -0.014 | -0.033 | 0.253 | 0.137 | -0.142 | 0.082
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0002 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 8.737 | -0.126 | -0.01 | 0.253 | -5.937 | -1.187 | 0.53
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.000125 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 18.273 | -0.247 | 0.087 | 0.253 | -19.21 | -2.263 | 1.085 | 摸底
case00 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0001 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 26.21 | -0.317 | 0.108 | 0.253 | -32.219 | -2.957 | 1.399 | 摸底
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 1 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.033 | -0.059 | -0.014 | -0.059 | -0.061 | -0.116 | -0.074
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.5 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.034 | -0.059 | -0.014 | -0.059 | -0.06 | -0.114 | -0.074
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.1 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.035 | -0.059 | -0.015 | -0.059 | -0.057 | -0.101 | -0.074
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.038 | -0.059 | -0.015 | -0.059 | -0.053 | -0.084 | -0.073
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.005 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.067 | -0.064 | -0.03 | -0.059 | -0.103 | 0.216 | -0.067
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0025 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.173 | -0.069 | -0.019 | -0.059 | -0.421 | 0.548 | -0.058
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.958 | -0.077 | -0.031 | -0.059 | -3.042 | 1.554 | -0.043
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -3.787 | -0.105 | -0.018 | -0.059 | -13.032 | 3.218 | 0.003
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0002 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -18.627 | -0.228 | -0.027 | -0.059 | -88.251 | 8.169 | 0.194
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.000125 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -50.747 | -0.349 | 0.034 | -0.059 | -250.205 | 13.135 | 0.281 | 摸底
case00 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0001 | -0.8 | 0 | 0 | 2 | 0 | 0 | -0.028 | -103.553 | -0.382 | -0.036 | -0.059 | -445.202 | 16.52 | 0.267 | 摸底
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 1 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.01 | 0.029 | 0.002 | 0.253 | 0.256 | 0.522 | -0.168
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.5 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.011 | 0.029 | 0.002 | 0.253 | 0.257 | 0.522 | -0.168
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.1 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.018 | 0.028 | 0.002 | 0.253 | 0.262 | 0.519 | -0.167
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.028 | 0.028 | 0.002 | 0.253 | 0.269 | 0.516 | -0.166
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.005 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.207 | 0.024 | -0.001 | 0.253 | 0.384 | 0.455 | -0.143
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0025 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.423 | 0.019 | -0.004 | 0.253 | 0.482 | 0.388 | -0.118
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 1.15 | 0.011 | -0.019 | 0.253 | 0.589 | 0.194 | -0.047
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 2.62 | -0.013 | -0.035 | 0.253 | 0.137 | -0.141 | 0.078
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0002 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 8.679 | -0.126 | 0.001 | 0.253 | -5.934 | -1.187 | 0.541
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.000125 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 19.449 | -0.267 | 0.069 | 0.253 | -19.23 | -2.283 | 1.067 | 摸底
case00 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0001 | -0.8 | 0 | 0 | 2 | 0 | 0 | 0.005 | 27.227 | -0.332 | 0.088 | 0.253 | -32.237 | -2.973 | 1.378 | 摸底

## case01 test report
dynamic_test | ADC_samp_freq | fund_wave_freq | ADC_OSR | scan_freq | PGA | fund_wave_amp_vol | fund_wave_amp_current | power_factor | vol_dc_bias | current_dc_bias | harmonic_wave_oder | amp_harmonic_vol | amp_harmonic_current | full_RMSU_err (RMSU电压有效值误差(‰)) | full_RMSI_err (RMSI电流有效值误差(‰)) | full_P_err (P有功功率误差(‰)) | full_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | 备注
case01 | 819200 | 50 | 128 | 49 | 1 | 0.4 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.034 | 0.039 | 0.078 |  | -0.397 | -0.393 | -0.785
case01 | 819200 | 50 | 128 | 49 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.034 | -0.021 | -0.024 |  | -0.397 | -0.438 | -0.871
case01 | 819200 | 50 | 128 | 49 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.034 | -1.644 | -2.576 |  | -0.397 | -2.398 | -3.023
case01 | 819200 | 50 | 128 | 49 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.034 | -33.016 | -26.5 |  | -0.397 | -103.718 | -23.198 |  | 摸底
case01 | 819200 | 50 | 128 | 51 | 1 | 0.4 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.116 | 0.114 | 0.235 |  | 0.498 | 0.495 | 0.998
case01 | 819200 | 50 | 128 | 51 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.116 | 0.145 | 0.225 |  | 0.498 | 0.541 | 1.004
case01 | 819200 | 50 | 128 | 51 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.116 | 0.842 | -0.028 |  | 0.498 | 0.892 | 1.163
case01 | 819200 | 50 | 128 | 51 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.116 | -4.333 | -2.385 |  | 0.498 | -76.732 | 2.652 |  | 摸底
case01 | 819200 | 50 | 256 | 49 | 1 | 0.4 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.096 | 0.095 | 0.206 |  | -0.611 | -0.613 | -1.208
case01 | 819200 | 50 | 256 | 49 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.096 | 0.116 | 0.253 |  | -0.611 | -0.595 | -1.164
case01 | 819200 | 50 | 256 | 49 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.096 | 0.717 | 1.457 |  | -0.611 | -0.228 | -0.044
case01 | 819200 | 50 | 256 | 49 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.096 | 9.799 | 12.746 |  | -0.611 | -5.579 | 10.466 |  | 摸底
case01 | 819200 | 50 | 256 | 51 | 1 | 0.4 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.03 | -0.03 | -0.044 |  | 1.049 | 1.049 | 2.115
case01 | 819200 | 50 | 256 | 51 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.03 | -0.03 | -0.019 |  | 1.049 | 1.042 | 2.133
case01 | 819200 | 50 | 256 | 51 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.03 | 0.005 | 0.626 |  | 1.049 | 0.758 | 2.603
case01 | 819200 | 50 | 256 | 51 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.03 | 4.169 | 6.675 |  | 1.049 | -10.638 | 7.006 |  | 摸底
case01 | 983040 | 60 | 128 | 59 | 1 | 0.4 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.078 | 0.083 | 0.166 |  | -0.325 | -0.322 | -0.642
case01 | 983040 | 60 | 128 | 59 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.078 | 0.028 | 0.069 |  | -0.325 | -0.364 | -0.725
case01 | 983040 | 60 | 128 | 59 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.078 | -1.452 | -2.36 |  | -0.325 | -2.259 | -2.811
case01 | 983040 | 60 | 128 | 59 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.078 | -32.984 | -25.141 |  | -0.325 | -103.019 | -22.377 |  | 摸底
case01 | 983040 | 60 | 128 | 61 | 1 | 0.4 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.097 | 0.094 | 0.196 |  | 0.429 | 0.425 | 0.859
case01 | 983040 | 60 | 128 | 61 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.097 | 0.128 | 0.189 |  | 0.429 | 0.475 | 0.869
case01 | 983040 | 60 | 128 | 61 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.097 | 0.96 | 0.005 |  | 0.429 | 0.929 | 1.13
case01 | 983040 | 60 | 128 | 61 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.097 | -11.938 | -1.721 |  | 0.429 | -75.603 | 3.581 |  | 摸底
case01 | 983040 | 60 | 256 | 59 | 1 | 0.4 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.058 | -0.06 | -0.103 |  | -0.606 | -0.608 | -1.198
case01 | 983040 | 60 | 256 | 59 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.058 | -0.026 | -0.042 |  | -0.606 | -0.578 | -1.143
case01 | 983040 | 60 | 256 | 59 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.058 | 0.91 | 1.496 |  | -0.606 | 0.085 | 0.275
case01 | 983040 | 60 | 256 | 59 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.058 | 13.418 | 15.922 |  | -0.606 | -2.459 | 13.569 |  | 摸底
case01 | 983040 | 60 | 256 | 61 | 1 | 0.4 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.086 | -0.086 | -0.156 |  | 0.906 | 0.907 | 1.829
case01 | 983040 | 60 | 256 | 61 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.086 | -0.084 | -0.128 |  | 0.906 | 0.903 | 1.852
case01 | 983040 | 60 | 256 | 61 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.086 | 0.011 | 0.576 |  | 0.906 | 0.732 | 2.434
case01 | 983040 | 60 | 256 | 61 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.086 | 4.205 | 7.198 |  | 0.906 | -9.598 | 7.905 |  | 摸底
case01 | 819200 | 50 | 128 | 49 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.034 | 0.075 | 0.672 | -0.038 | -0.397 | -0.492 | -0.35 | -1.028
case01 | 819200 | 50 | 128 | 49 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.034 | 0.05 | 0.47 | -0.079 | -0.397 | -0.515 | -0.521 | -1.082
case01 | 819200 | 50 | 128 | 49 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.034 | -0.707 | -4.638 | -1.108 | -0.397 | -1.923 | -4.827 | -2.427
case01 | 819200 | 50 | 128 | 49 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.034 | -28.442 | -52.607 | -10.694 | -0.397 | -97.64 | -45.275 | -15 | 摸底
case01 | 819200 | 50 | 128 | 51 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.116 | -0.189 | -0.11 | -0.023 | 0.498 | 0.196 | 0.598 | 0.761
case01 | 819200 | 50 | 128 | 51 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.116 | -0.183 | -0.131 | -0.079 | 0.498 | 0.202 | 0.611 | 0.685
case01 | 819200 | 50 | 128 | 51 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.116 | -0.138 | -0.638 | -1.483 | 0.498 | -0.456 | 0.924 | -1.207
case01 | 819200 | 50 | 128 | 51 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.116 | -16.789 | -5.42 | -14.568 | 0.498 | -88.363 | 3.878 | -18.942 | 摸底
case01 | 819200 | 50 | 256 | 49 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.096 | -0.13 | -0.138 | -0.014 | -0.611 | -0.836 | -1.555 | -2.04
case01 | 819200 | 50 | 256 | 49 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.096 | -0.091 | -0.043 | 0.028 | -0.611 | -0.802 | -1.466 | -2.001
case01 | 819200 | 50 | 256 | 49 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.096 | 0.94 | 2.361 | 1.097 | -0.611 | -0.03 | 0.77 | -1.006
case01 | 819200 | 50 | 256 | 49 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.096 | 14.395 | 24.831 | 11.17 | -0.611 | -1.52 | 21.662 | 8.382 | 摸底
case01 | 819200 | 50 | 256 | 51 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.03 | 0.14 | 0.365 | 0.026 | 1.049 | 1.256 | 2.452 | 1.546
case01 | 819200 | 50 | 256 | 51 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.03 | 0.138 | 0.416 | 0.029 | 1.049 | 1.249 | 2.489 | 1.548
case01 | 819200 | 50 | 256 | 51 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.03 | 0.139 | 1.701 | 0.101 | 1.049 | 0.992 | 3.423 | 1.625
case01 | 819200 | 50 | 256 | 51 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.03 | 3.923 | 13.592 | 0.85 | 1.049 | -10.127 | 12.019 | 2.427 | 摸底
case01 | 983040 | 60 | 128 | 59 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.078 | -0.043 | 0.372 | -0.036 | -0.325 | -0.512 | -0.61 | -0.873
case01 | 983040 | 60 | 128 | 59 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.078 | -0.048 | 0.18 | -0.056 | -0.325 | -0.522 | -0.776 | -0.91
case01 | 983040 | 60 | 128 | 59 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.078 | -0.219 | -4.684 | -0.552 | -0.325 | -1.594 | -4.953 | -1.859
case01 | 983040 | 60 | 128 | 59 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.078 | -18.744 | -50.348 | -5.145 | -0.325 | -93.696 | -44.174 | -10.697 | 摸底
case01 | 983040 | 60 | 128 | 61 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.097 | -0.088 | 0.217 | -0.023 | 0.429 | 0.227 | 0.821 | 0.637
case01 | 983040 | 60 | 128 | 61 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.097 | -0.073 | 0.202 | -0.064 | 0.429 | 0.25 | 0.841 | 0.578
case01 | 983040 | 60 | 128 | 61 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.097 | 0.201 | -0.17 | -1.067 | 0.429 | 0 | 1.361 | -0.891
case01 | 983040 | 60 | 128 | 61 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.097 | -17.962 | -3.677 | -10.383 | 0.429 | -83.756 | 6.239 | -14.646 | 摸底
case01 | 983040 | 60 | 256 | 59 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.058 | -0.037 | -0.395 | -0.008 | -0.606 | -0.588 | -1.495 | -1.729
case01 | 983040 | 60 | 256 | 59 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.058 | -0.01 | -0.274 | 0.012 | -0.606 | -0.564 | -1.383 | -1.711
case01 | 983040 | 60 | 256 | 59 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.058 | 0.735 | 2.799 | 0.523 | -0.606 | -0.038 | 1.447 | -1.238
case01 | 983040 | 60 | 256 | 59 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.058 | 11.445 | 31.526 | 5.362 | -0.606 | -3.824 | 27.891 | 3.253 | 摸底
case01 | 983040 | 60 | 256 | 61 | 1 | 0.4 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.086 | 0.057 | -0.186 | 0.024 | 0.906 | 1.013 | 1.695 | 1.291
case01 | 983040 | 60 | 256 | 61 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.086 | 0.057 | -0.131 | 0.027 | 0.906 | 1.008 | 1.741 | 1.292
case01 | 983040 | 60 | 256 | 61 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.086 | 0.126 | 1.275 | 0.126 | 0.906 | 0.788 | 2.901 | 1.317
case01 | 983040 | 60 | 256 | 61 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.086 | 4.212 | 14.31 | 1.13 | 0.906 | -9.969 | 13.629 | 1.633 | 摸底

## case02 test report
dynamic_test | ADC_samp_freq | fund_wave_freq | ADC_OSR | scan_freq | PGA | fund_wave_amp_vol | fund_wave_amp_current | power_factor | vol_dc_bias | current_dc_bias | harmonic_wave_oder | amp_harmonic_vol | amp_harmonic_current | full_RMSU_err (RMSU电压有效值误差(‰)) | full_RMSI_err (RMSI电流有效值误差(‰)) | full_P_err (P有功功率误差(‰)) | full_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | 备注
case02 | 819200 | 50 | 128 | 45 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.27 | -0.235 | -0.529 |  | -2.542 | -2.525 | -5.081
case02 | 819200 | 50 | 128 | 45 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.27 | 0.635 | -0.602 |  | -2.542 | -2.888 | -5.633
case02 | 819200 | 50 | 128 | 45 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.27 | 1.282 | -0.678 |  | -2.542 | -4.972 | -6.21
case02 | 819200 | 50 | 128 | 45 | 1 | 0.4 | 0.0004 | 1 | 0 | 0 | 2 | 0 | 0 | -0.27 | -0.273 | -0.904 |  | -2.542 | -21.818 | -7.937
case02 | 819200 | 50 | 128 | 45 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.27 | -13.145 | -1.283 |  | -2.542 | -87.877 | -10.814 |  | 摸底
case02 | 819200 | 50 | 128 | 48 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.277 | 0.301 | 0.542 |  | -0.954 | -0.941 | -1.929
case02 | 819200 | 50 | 128 | 48 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.277 | 0.826 | 0.146 |  | -0.954 | -1.401 | -2.583
case02 | 819200 | 50 | 128 | 48 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | 0.277 | 0.922 | -0.265 |  | -0.954 | -3.584 | -3.263
case02 | 819200 | 50 | 128 | 48 | 1 | 0.4 | 0.0004 | 1 | 0 | 0 | 2 | 0 | 0 | 0.277 | -0.491 | -1.499 |  | -0.954 | -20.729 | -5.304
case02 | 819200 | 50 | 128 | 48 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.277 | -11.639 | -3.554 |  | -0.954 | -87.316 | -8.713 |  | 摸底
case02 | 819200 | 50 | 128 | 52 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.139 | -0.156 | -0.331 |  | 0.481 | 0.45 | 0.896
case02 | 819200 | 50 | 128 | 52 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.139 | -0.722 | -1.831 |  | 0.481 | -1.228 | -0.975
case02 | 819200 | 50 | 128 | 52 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.139 | -1.738 | -3.393 |  | 0.481 | -4.686 | -2.925
case02 | 819200 | 50 | 128 | 52 | 1 | 0.4 | 0.0004 | 1 | 0 | 0 | 2 | 0 | 0 | -0.139 | -6.833 | -8.073 |  | 0.481 | -25.756 | -8.768
case02 | 819200 | 50 | 128 | 52 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.139 | -26.297 | -15.87 |  | 0.481 | -99.697 | -18.509 |  | 摸底
case02 | 819200 | 50 | 128 | 55 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.365 | 0.333 | 0.668 |  | -1.568 | -1.578 | -3.173
case02 | 819200 | 50 | 128 | 55 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.365 | -0.622 | -1.236 |  | -1.568 | -2.68 | -4.459
case02 | 819200 | 50 | 128 | 55 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | 0.365 | -2.101 | -3.218 |  | -1.568 | -5.542 | -5.798
case02 | 819200 | 50 | 128 | 55 | 1 | 0.4 | 0.0004 | 1 | 0 | 0 | 2 | 0 | 0 | 0.365 | -8.002 | -9.16 |  | -1.568 | -24.829 | -9.807
case02 | 819200 | 50 | 128 | 55 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.365 | -27.226 | -19.067 |  | -1.568 | -95.645 | -16.506 |  | 摸底
case02 | 819200 | 50 | 256 | 45 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.493 | 0.486 | 1.027 |  | -2.73 | -2.735 | -5.41
case02 | 819200 | 50 | 256 | 45 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.493 | 0.355 | 1.498 |  | -2.73 | -2.955 | -4.878
case02 | 819200 | 50 | 256 | 45 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | 0.493 | 0.337 | 1.988 |  | -2.73 | -3.374 | -4.324
case02 | 819200 | 50 | 256 | 45 | 1 | 0.4 | 0.0004 | 1 | 0 | 0 | 2 | 0 | 0 | 0.493 | 0.801 | 3.459 |  | -2.73 | -5.784 | -2.663
case02 | 819200 | 50 | 256 | 45 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.493 | 3.187 | 5.908 |  | -2.73 | -13.688 | 0.106 |  | 摸底
case02 | 819200 | 50 | 256 | 48 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.009 | 0.02 | 0.07 |  | -1.567 | -1.557 | -3.08
case02 | 819200 | 50 | 256 | 48 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.009 | 0.387 | 1.039 |  | -1.567 | -1.359 | -2.128
case02 | 819200 | 50 | 256 | 48 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | 0.009 | 0.835 | 2.048 |  | -1.567 | -1.344 | -1.136
case02 | 819200 | 50 | 256 | 48 | 1 | 0.4 | 0.0004 | 1 | 0 | 0 | 2 | 0 | 0 | 0.009 | 2.866 | 5.075 |  | -1.567 | -2.468 | 1.838
case02 | 819200 | 50 | 256 | 48 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.009 | 7.173 | 10.13 |  | -1.567 | -8.252 | 6.807 |  | 摸底
case02 | 819200 | 50 | 256 | 52 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.049 | 0.058 | 0.148 |  | 1.7 | 1.72 | 3.465
case02 | 819200 | 50 | 256 | 52 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.049 | 0.367 | 1.065 |  | 1.7 | 2.16 | 4.659
case02 | 819200 | 50 | 256 | 52 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | 0.049 | 0.768 | 2.019 |  | 1.7 | 2.428 | 5.903
case02 | 819200 | 50 | 256 | 52 | 1 | 0.4 | 0.0004 | 1 | 0 | 0 | 2 | 0 | 0 | 0.049 | 2.413 | 4.887 |  | 1.7 | 2.072 | 9.641
case02 | 819200 | 50 | 256 | 52 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.049 | 7.441 | 9.661 |  | 1.7 | -2.423 | 15.864 |  | 摸底
case02 | 819200 | 50 | 256 | 55 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.026 | 0.011 | 0.081 |  | -2.287 | -2.307 | -4.544
case02 | 819200 | 50 | 256 | 55 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.026 | -0.329 | 0.343 |  | -2.287 | -2.928 | -4.41
case02 | 819200 | 50 | 256 | 55 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | 0.026 | -0.604 | 0.615 |  | -2.287 | -3.765 | -4.272
case02 | 819200 | 50 | 256 | 55 | 1 | 0.4 | 0.0004 | 1 | 0 | 0 | 2 | 0 | 0 | 0.026 | -0.915 | 1.441 |  | -2.287 | -7.425 | -3.848
case02 | 819200 | 50 | 256 | 55 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.026 | 0.693 | 2.826 |  | -2.287 | -17.427 | -3.13 |  | 摸底
case02 | 983040 | 60 | 128 | 55 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.074 | -0.074 | -0.177 |  | -1.613 | -1.621 | -3.26
case02 | 983040 | 60 | 128 | 55 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.074 | -0.15 | -1.208 |  | -1.613 | -2.642 | -4.475
case02 | 983040 | 60 | 128 | 55 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.074 | -0.748 | -2.28 |  | -1.613 | -5.404 | -5.739
case02 | 983040 | 60 | 128 | 55 | 1 | 0.4 | 0.0004 | 1 | 0 | 0 | 2 | 0 | 0 | -0.074 | -4.174 | -5.508 |  | -1.613 | -24.311 | -9.54
case02 | 983040 | 60 | 128 | 55 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.074 | -23.323 | -10.867 |  | -1.613 | -94.033 | -15.859 |  | 摸底
case02 | 983040 | 60 | 128 | 58 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.36 | -0.371 | -0.767 |  | -1.263 | -1.275 | -2.572
case02 | 983040 | 60 | 128 | 58 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.36 | -0.717 | -2.095 |  | -1.263 | -2.432 | -3.922
case02 | 983040 | 60 | 128 | 58 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.36 | -1.654 | -3.477 |  | -1.263 | -5.338 | -5.327
case02 | 983040 | 60 | 128 | 58 | 1 | 0.4 | 0.0004 | 1 | 0 | 0 | 2 | 0 | 0 | -0.36 | -5.435 | -7.62 |  | -1.263 | -24.692 | -9.54
case02 | 983040 | 60 | 128 | 58 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.36 | -24.418 | -14.534 |  | -1.263 | -95.339 | -16.571 |  | 摸底
case02 | 983040 | 60 | 128 | 62 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.423 | -0.424 | -0.883 |  | 0.164 | 0.157 | 0.285
case02 | 983040 | 60 | 128 | 62 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.423 | -0.576 | -1.948 |  | 0.164 | -0.846 | -0.91
case02 | 983040 | 60 | 128 | 62 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.423 | -0.915 | -3.056 |  | 0.164 | -3.596 | -2.154
case02 | 983040 | 60 | 128 | 62 | 1 | 0.4 | 0.0004 | 1 | 0 | 0 | 2 | 0 | 0 | -0.423 | -5.127 | -6.378 |  | 0.164 | -22.485 | -5.882
case02 | 983040 | 60 | 128 | 62 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.423 | -16.647 | -11.926 |  | 0.164 | -92.322 | -12.115 |  | 摸底
case02 | 983040 | 60 | 128 | 65 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.043 | -0.026 | -0.103 |  | -0.338 | -0.326 | -0.696
case02 | 983040 | 60 | 128 | 65 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.043 | 0.422 | -0.686 |  | -0.338 | -0.817 | -1.381
case02 | 983040 | 60 | 128 | 65 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.043 | 0.403 | -1.295 |  | -0.338 | -3.034 | -2.095
case02 | 983040 | 60 | 128 | 65 | 1 | 0.4 | 0.0004 | 1 | 0 | 0 | 2 | 0 | 0 | -0.043 | -1.75 | -3.108 |  | -0.338 | -20.291 | -4.23
case02 | 983040 | 60 | 128 | 65 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.043 | -9.436 | -6.151 |  | -0.338 | -87.095 | -7.804 |  | 摸底
case02 | 983040 | 60 | 256 | 55 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.043 | -0.019 | -0.016 |  | -2.854 | -2.826 | -5.626
case02 | 983040 | 60 | 256 | 55 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.043 | 0.71 | 1.317 |  | -2.854 | -2.148 | -4.197
case02 | 983040 | 60 | 256 | 55 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.043 | 1.549 | 2.706 |  | -2.854 | -1.631 | -2.71
case02 | 983040 | 60 | 256 | 55 | 1 | 0.4 | 0.0004 | 1 | 0 | 0 | 2 | 0 | 0 | -0.043 | 4.593 | 6.869 |  | -2.854 | -1.239 | 1.752
case02 | 983040 | 60 | 256 | 55 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.043 | 11.167 | 13.818 |  | -2.854 | -4.434 | 9.197 |  | 摸底
case02 | 983040 | 60 | 256 | 58 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.013 | -0.051 | -0.023 |  | -1.333 | -1.37 | -2.659
case02 | 983040 | 60 | 256 | 58 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.013 | -1.053 | -0.419 |  | -1.333 | -2.46 | -2.993
case02 | 983040 | 60 | 256 | 58 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.013 | -2.012 | -0.831 |  | -1.333 | -3.788 | -3.34
case02 | 983040 | 60 | 256 | 58 | 1 | 0.4 | 0.0004 | 1 | 0 | 0 | 2 | 0 | 0 | -0.013 | -4.408 | -2.069 |  | -1.333 | -8.967 | -4.382
case02 | 983040 | 60 | 256 | 58 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.013 | -7.068 | -4.136 |  | -1.333 | -21.657 | -6.123 |  | 摸底
case02 | 983040 | 60 | 256 | 62 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.099 | -0.097 | -0.155 |  | 1.441 | 1.447 | 2.931
case02 | 983040 | 60 | 256 | 62 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.099 | -0.018 | 0.538 |  | 1.441 | 1.521 | 3.759
case02 | 983040 | 60 | 256 | 62 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.099 | 0.167 | 1.26 |  | 1.441 | 1.411 | 4.622
case02 | 983040 | 60 | 256 | 62 | 1 | 0.4 | 0.0004 | 1 | 0 | 0 | 2 | 0 | 0 | -0.099 | 1.09 | 3.429 |  | 1.441 | -0.072 | 7.215
case02 | 983040 | 60 | 256 | 62 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.099 | 4.829 | 7.038 |  | 1.441 | -6.428 | 11.532 |  | 摸底
case02 | 983040 | 60 | 256 | 65 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.019 | 0.024 | 0.087 |  | -0.043 | -0.038 | -0.038
case02 | 983040 | 60 | 256 | 65 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.019 | 0.229 | 0.886 |  | -0.043 | -0.002 | 0.748
case02 | 983040 | 60 | 256 | 65 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | 0.019 | 0.51 | 1.72 |  | -0.043 | -0.149 | 1.569
case02 | 983040 | 60 | 256 | 65 | 1 | 0.4 | 0.0004 | 1 | 0 | 0 | 2 | 0 | 0 | 0.019 | 1.987 | 4.222 |  | -0.043 | -1.73 | 4.031
case02 | 983040 | 60 | 256 | 65 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.019 | 6.168 | 8.395 |  | -0.043 | -8.188 | 8.139 |  | 摸底
case02 | 819200 | 50 | 128 | 45 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.27 | -0.729 | -3.61 | -0.206 | -2.542 | -1.578 | -6.315 | -3.45
case02 | 819200 | 50 | 128 | 45 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.27 | -0.076 | -3.753 | -1.142 | -2.542 | -1.895 | -7.419 | -4.245
case02 | 819200 | 50 | 128 | 45 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.27 | 0.26 | -3.925 | -2.166 | -2.542 | -3.922 | -8.587 | -5.066
case02 | 819200 | 50 | 128 | 45 | 1 | 0.4 | 0.0004 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.27 | -1.959 | -4.349 | -5.136 | -2.542 | -20.555 | -12.029 | -7.555
case02 | 819200 | 50 | 128 | 45 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.27 | -12.83 | -5.213 | -10.093 | -2.542 | -85.952 | -17.893 | -11.643 | 摸底
case02 | 819200 | 50 | 128 | 48 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.277 | -0.636 | -1.138 | -0.139 | -0.954 | -1.496 | -3.937 | -1.99
case02 | 819200 | 50 | 128 | 48 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.277 | -1.329 | -1.929 | -2.654 | -0.954 | -3.031 | -5.244 | -4.386
case02 | 819200 | 50 | 128 | 48 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.277 | -2.706 | -2.758 | -5.285 | -0.954 | -6.334 | -6.615 | -6.878
case02 | 819200 | 50 | 128 | 48 | 1 | 0.4 | 0.0004 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.277 | -10.026 | -5.241 | -13.13 | -0.954 | -26.91 | -10.725 | -14.359
case02 | 819200 | 50 | 128 | 48 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.277 | -26.287 | -9.4 | -26.38 | -0.954 | -99.797 | -17.546 | -26.818 | 摸底
case02 | 819200 | 50 | 128 | 52 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.139 | 0.4 | 1.244 | -0.126 | 0.481 | 1.033 | 2.435 | 1.169
case02 | 819200 | 50 | 128 | 52 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.139 | -0.413 | -1.754 | -2.359 | 0.481 | -0.943 | -1.307 | -0.781
case02 | 819200 | 50 | 128 | 52 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.139 | -1.732 | -4.873 | -4.667 | 0.481 | -4.712 | -5.199 | -2.811
case02 | 819200 | 50 | 128 | 52 | 1 | 0.4 | 0.0004 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.139 | -8.918 | -14.286 | -11.56 | 0.481 | -26.752 | -16.947 | -8.877
case02 | 819200 | 50 | 128 | 52 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.139 | -28.714 | -29.94 | -23.174 | 0.481 | -102.557 | -36.421 | -19.015 | 摸底
case02 | 819200 | 50 | 128 | 55 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.365 | 0.095 | 2.101 | -0.151 | -1.568 | -1.245 | -1.18 | -3.413
case02 | 819200 | 50 | 128 | 55 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.365 | -0.549 | -1.704 | -1.45 | -1.568 | -2.171 | -3.747 | -4.416
case02 | 819200 | 50 | 128 | 55 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.365 | -1.557 | -5.657 | -2.792 | -1.568 | -4.845 | -6.413 | -5.463
case02 | 819200 | 50 | 128 | 55 | 1 | 0.4 | 0.0004 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.365 | -7.107 | -17.628 | -6.706 | -1.568 | -23.525 | -14.529 | -8.557
case02 | 819200 | 50 | 128 | 55 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.365 | -22.878 | -37.446 | -13.518 | -1.568 | -93.013 | -27.854 | -13.776 | 摸底
case02 | 819200 | 50 | 256 | 45 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.493 | -0.578 | 0.246 | -0.273 | -2.73 | -3.687 | -7.107 | -6.624
case02 | 819200 | 50 | 256 | 45 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.493 | -1.657 | 1.187 | -1.598 | -2.73 | -5.026 | -6.043 | -8.111
case02 | 819200 | 50 | 256 | 45 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.493 | -2.703 | 2.163 | -2.976 | -2.73 | -6.612 | -4.94 | -9.656
case02 | 819200 | 50 | 256 | 45 | 1 | 0.4 | 0.0004 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.493 | -5.352 | 5.084 | -7.1 | -2.73 | -12.559 | -1.631 | -14.288
case02 | 819200 | 50 | 256 | 45 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.493 | -7.804 | 9.967 | -13.994 | -2.73 | -26.546 | 3.882 | -22.021 | 摸底
case02 | 819200 | 50 | 256 | 48 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.009 | -0.025 | 0.21 | -0.096 | -1.567 | -1.61 | -2.951 | -3.816
case02 | 819200 | 50 | 256 | 48 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.009 | -0.903 | 2.144 | -1.371 | -1.567 | -2.616 | -1.05 | -5.079
case02 | 819200 | 50 | 256 | 48 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.009 | -1.749 | 4.155 | -2.698 | -1.567 | -3.855 | 0.926 | -6.392
case02 | 819200 | 50 | 256 | 48 | 1 | 0.4 | 0.0004 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.009 | -3.776 | 10.24 | -6.701 | -1.567 | -8.769 | 6.912 | -10.36
case02 | 819200 | 50 | 256 | 48 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.009 | -5.428 | 20.189 | -13.291 | -1.567 | -20.987 | 16.661 | -16.873 | 摸底
case02 | 819200 | 50 | 256 | 52 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.049 | 0.025 | 0.23 | 0.051 | 1.7 | 1.703 | 3.598 | 2.612
case02 | 819200 | 50 | 256 | 52 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.049 | 0.25 | 2.061 | 0.259 | 1.7 | 1.872 | 5.985 | 2.725
case02 | 819200 | 50 | 256 | 52 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.049 | 0.558 | 3.974 | 0.474 | 1.7 | 1.86 | 8.476 | 2.839
case02 | 819200 | 50 | 256 | 52 | 1 | 0.4 | 0.0004 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.049 | 2.069 | 9.747 | 1.102 | 1.7 | 0.666 | 15.989 | 3.165
case02 | 819200 | 50 | 256 | 52 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.049 | 6.07 | 19.044 | 2.305 | 1.7 | -5.148 | 28.181 | 3.871 | 摸底
case02 | 819200 | 50 | 256 | 55 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.026 | -0.331 | -1.282 | 0.061 | -2.287 | -2.968 | -5.146 | -6.123
case02 | 819200 | 50 | 256 | 55 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.026 | -0.224 | -0.757 | 0.66 | -2.287 | -2.833 | -4.877 | -5.202
case02 | 819200 | 50 | 256 | 55 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.026 | -0.012 | -0.208 | 1.283 | -2.287 | -2.878 | -4.595 | -4.242
case02 | 819200 | 50 | 256 | 55 | 1 | 0.4 | 0.0004 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.026 | 1.186 | 1.492 | 3.131 | -2.287 | -4.158 | -3.7 | -1.386
case02 | 819200 | 50 | 256 | 55 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.026 | 4.962 | 3.909 | 6.393 | -2.287 | -10.08 | -2.621 | 3.552 | 摸底
case02 | 983040 | 60 | 128 | 55 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.074 | -0.096 | -0.226 | -0.203 | -1.613 | -1.639 | -3.316 | -3.277
case02 | 983040 | 60 | 128 | 55 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.074 | -0.736 | -2.287 | -2.041 | -1.613 | -3.129 | -5.746 | -5.017
case02 | 983040 | 60 | 128 | 55 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.074 | -1.554 | -4.451 | -3.967 | -1.613 | -6.387 | -8.291 | -6.825
case02 | 983040 | 60 | 128 | 55 | 1 | 0.4 | 0.0004 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.074 | -6.852 | -10.883 | -9.599 | -1.613 | -26.848 | -15.877 | -12.268
case02 | 983040 | 60 | 128 | 55 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.074 | -29.377 | -21.72 | -19.079 | -1.613 | -99.649 | -28.612 | -21.284 | 摸底
case02 | 983040 | 60 | 128 | 58 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.36 | -0.202 | -1.916 | -0.143 | -1.263 | -0.757 | -3.034 | -1.718
case02 | 983040 | 60 | 128 | 58 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.36 | -1.611 | -4.574 | -2.782 | -1.263 | -2.685 | -5.736 | -4.187
case02 | 983040 | 60 | 128 | 58 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.36 | -3.577 | -7.348 | -5.555 | -1.263 | -6.4 | -8.556 | -6.755
case02 | 983040 | 60 | 128 | 58 | 1 | 0.4 | 0.0004 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.36 | -11.883 | -15.653 | -13.888 | -1.263 | -28.252 | -17.014 | -14.463
case02 | 983040 | 60 | 128 | 58 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.36 | -30.929 | -29.535 | -27.814 | -1.263 | -103.59 | -31.098 | -27.291 | 摸底
case02 | 983040 | 60 | 128 | 62 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.423 | 0.154 | -0.859 | -0.127 | 0.164 | 0.725 | 0.309 | 1.044
case02 | 983040 | 60 | 128 | 62 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.423 | -0.636 | -2.991 | -2.531 | 0.164 | -1.248 | -2.083 | -1.421
case02 | 983040 | 60 | 128 | 62 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.423 | -1.817 | -5.204 | -5.019 | 0.164 | -5.015 | -4.567 | -3.991
case02 | 983040 | 60 | 128 | 62 | 1 | 0.4 | 0.0004 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.423 | -7.769 | -11.889 | -12.541 | 0.164 | -27.044 | -12.081 | -11.679
case02 | 983040 | 60 | 128 | 62 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.423 | -27.963 | -23.026 | -25.021 | 0.164 | -102.78 | -24.528 | -24.491 | 摸底
case02 | 983040 | 60 | 128 | 65 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.043 | -0.061 | -0.133 | -0.155 | -0.338 | -0.343 | -0.727 | -0.707
case02 | 983040 | 60 | 128 | 65 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.043 | -0.495 | -1.295 | -2.124 | -0.338 | -1.253 | -2.091 | -1.711
case02 | 983040 | 60 | 128 | 65 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.043 | -1.429 | -2.5 | -4.202 | -0.338 | -3.91 | -3.51 | -2.757
case02 | 983040 | 60 | 128 | 65 | 1 | 0.4 | 0.0004 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.043 | -7.548 | -6.191 | -10.346 | -0.338 | -22.531 | -7.85 | -5.85
case02 | 983040 | 60 | 128 | 65 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.043 | -19.373 | -12.277 | -20.547 | -0.338 | -91.962 | -14.944 | -11.077 | 摸底
case02 | 983040 | 60 | 256 | 55 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.043 | 0.02 | 0.052 | -0.098 | -2.854 | -2.791 | -5.554 | -6.141
case02 | 983040 | 60 | 256 | 55 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.043 | 1.69 | 2.72 | 1.725 | -2.854 | -1.249 | -2.695 | -4.289
case02 | 983040 | 60 | 256 | 55 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.043 | 3.49 | 5.49 | 3.628 | -2.854 | 0.167 | 0.272 | -2.355
case02 | 983040 | 60 | 256 | 55 | 1 | 0.4 | 0.0004 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.043 | 9.48 | 13.82 | 9.334 | -2.854 | 3.247 | 9.203 | 3.44
case02 | 983040 | 60 | 256 | 55 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.043 | 20.791 | 27.659 | 18.851 | -2.854 | 4.506 | 24.02 | 13.117 | 摸底
case02 | 983040 | 60 | 256 | 58 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.013 | -0.161 | -0.514 | -0.06 | -1.333 | -1.455 | -3.142 | -3.247
case02 | 983040 | 60 | 256 | 58 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.013 | -1.124 | -1.311 | -0.546 | -1.333 | -2.51 | -3.814 | -3.73
case02 | 983040 | 60 | 256 | 58 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.013 | -2.049 | -2.141 | -1.051 | -1.333 | -3.8 | -4.513 | -4.233
case02 | 983040 | 60 | 256 | 58 | 1 | 0.4 | 0.0004 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.013 | -4.276 | -4.587 | -2.58 | -1.333 | -8.86 | -6.561 | -5.757
case02 | 983040 | 60 | 256 | 58 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.013 | -6.134 | -8.873 | -5.053 | -1.333 | -21.321 | -10.218 | -8.203 | 摸底
case02 | 983040 | 60 | 256 | 62 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.099 | 0.098 | -0.078 | 0.056 | 1.441 | 1.593 | 2.834 | 2.392
case02 | 983040 | 60 | 256 | 62 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.099 | 0.336 | 1.306 | 0.546 | 1.441 | 1.945 | 4.489 | 3.108
case02 | 983040 | 60 | 256 | 62 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.099 | 0.672 | 2.751 | 1.054 | 1.441 | 2.12 | 6.216 | 3.853
case02 | 983040 | 60 | 256 | 62 | 1 | 0.4 | 0.0004 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.099 | 2.207 | 7.134 | 2.56 | 1.441 | 1.48 | 11.445 | 6.066
case02 | 983040 | 60 | 256 | 62 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.099 | 6.061 | 14.079 | 5.222 | 1.441 | -3.458 | 19.798 | 9.911 | 摸底
case02 | 983040 | 60 | 256 | 65 | 1 | 0.4 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.019 | -0.006 | 0.113 | 0.011 | -0.043 | -0.076 | -0.014 | -0.972
case02 | 983040 | 60 | 256 | 65 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.019 | -0.501 | 1.712 | -0.769 | -0.043 | -0.919 | 1.56 | -1.985
case02 | 983040 | 60 | 256 | 65 | 1 | 0.4 | 0.001 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.019 | -0.903 | 3.381 | -1.582 | -0.043 | -1.983 | 3.202 | -3.04
case02 | 983040 | 60 | 256 | 65 | 1 | 0.4 | 0.0004 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.019 | -1.628 | 8.437 | -4.04 | -0.043 | -6.336 | 8.174 | -6.226
case02 | 983040 | 60 | 256 | 65 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.019 | -1.039 | 16.47 | -7.966 | -0.043 | -17.472 | 16.076 | -11.369 | 摸底
case02 | 819200 | 50 | 128 | 45 | 1 | 0.4 | 0.05 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.27 | 0.119 |  | -0.2 | -2.542 | -0.869 |  | -3.445
case02 | 819200 | 50 | 128 | 45 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.27 | 0.355 |  | -1.037 | -2.542 | -1.395 |  | -4.133
case02 | 819200 | 50 | 128 | 45 | 1 | 0.4 | 0.001 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.27 | 0.162 |  | -1.854 | -2.542 | -3.643 |  | -4.85
case02 | 819200 | 50 | 128 | 45 | 1 | 0.4 | 0.0004 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.27 | -3.892 |  | -4.53 | -2.542 | -20.954 |  | -6.999
case02 | 819200 | 50 | 128 | 45 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.27 | -20.295 |  | -8.766 | -2.542 | -87.689 |  | -10.585 | 摸底
case02 | 819200 | 50 | 128 | 48 | 1 | 0.4 | 0.05 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.277 | -0.399 |  | -0.128 | -0.954 | -1.022 |  | -1.979
case02 | 819200 | 50 | 128 | 48 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.277 | -1.577 |  | -2.305 | -0.954 | -2.897 |  | -4.056
case02 | 819200 | 50 | 128 | 48 | 1 | 0.4 | 0.001 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.277 | -3.452 |  | -4.603 | -0.954 | -6.556 |  | -6.217
case02 | 819200 | 50 | 128 | 48 | 1 | 0.4 | 0.0004 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.277 | -10.89 |  | -11.437 | -0.954 | -28.229 |  | -12.703
case02 | 819200 | 50 | 128 | 48 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.277 | -34.774 |  | -22.764 | -0.954 | -103.251 |  | -23.516 | 摸底
case02 | 819200 | 50 | 128 | 52 | 1 | 0.4 | 0.05 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.139 | 0.048 |  | -0.117 | 0.481 | 0.702 |  | 1.178
case02 | 819200 | 50 | 128 | 52 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.139 | -0.626 |  | -2.047 | 0.481 | -0.977 |  | -0.511
case02 | 819200 | 50 | 128 | 52 | 1 | 0.4 | 0.001 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.139 | -1.831 |  | -4.021 | 0.481 | -4.437 |  | -2.27
case02 | 819200 | 50 | 128 | 52 | 1 | 0.4 | 0.0004 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.139 | -7.142 |  | -10.067 | 0.481 | -25.531 |  | -7.548
case02 | 819200 | 50 | 128 | 52 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.139 | -30.094 |  | -19.967 | 0.481 | -99.565 |  | -16.325 | 摸底
case02 | 819200 | 50 | 128 | 55 | 1 | 0.4 | 0.05 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.365 | -0.477 |  | -0.144 | -1.568 | -1.814 |  | -3.407
case02 | 819200 | 50 | 128 | 55 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.365 | -0.71 |  | -1.267 | -1.568 | -2.623 |  | -4.275
case02 | 819200 | 50 | 128 | 55 | 1 | 0.4 | 0.001 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.365 | -1.272 |  | -2.424 | -1.568 | -5.174 |  | -5.179
case02 | 819200 | 50 | 128 | 55 | 1 | 0.4 | 0.0004 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.365 | -4.634 |  | -5.94 | -1.568 | -23.466 |  | -7.888
case02 | 819200 | 50 | 128 | 55 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.365 | -27.389 |  | -11.774 | -1.568 | -92.188 |  | -12.374 | 摸底
case02 | 819200 | 50 | 256 | 45 | 1 | 0.4 | 0.05 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.493 | -0.66 |  | -0.27 | -2.73 | -3.423 |  | -6.62
case02 | 819200 | 50 | 256 | 45 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.493 | -1.81 |  | -1.418 | -2.73 | -4.883 |  | -7.908
case02 | 819200 | 50 | 256 | 45 | 1 | 0.4 | 0.001 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.493 | -2.895 |  | -2.612 | -2.73 | -6.596 |  | -9.248
case02 | 819200 | 50 | 256 | 45 | 1 | 0.4 | 0.0004 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.493 | -5.678 |  | -6.203 | -2.73 | -12.939 |  | -13.276
case02 | 819200 | 50 | 256 | 45 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.493 | -8.628 |  | -12.167 | -2.73 | -27.566 |  | -19.968 | 摸底
case02 | 819200 | 50 | 256 | 48 | 1 | 0.4 | 0.05 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.009 | -0.073 |  | -0.092 | -1.567 | -1.658 |  | -3.812
case02 | 819200 | 50 | 256 | 48 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.009 | -1.28 |  | -1.199 | -1.567 | -2.976 |  | -4.908
case02 | 819200 | 50 | 256 | 48 | 1 | 0.4 | 0.001 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.009 | -2.43 |  | -2.349 | -1.567 | -4.543 |  | -6.048
case02 | 819200 | 50 | 256 | 48 | 1 | 0.4 | 0.0004 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.009 | -5.457 |  | -5.805 | -1.567 | -10.442 |  | -9.469
case02 | 819200 | 50 | 256 | 48 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.009 | -8.901 |  | -11.557 | -1.567 | -24.346 |  | -15.168 | 摸底
case02 | 819200 | 50 | 256 | 52 | 1 | 0.4 | 0.05 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.049 | -0.005 |  | 0.05 | 1.7 | 1.662 |  | 2.612
case02 | 819200 | 50 | 256 | 52 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.049 | 0.095 |  | 0.229 | 1.7 | 1.562 |  | 2.707
case02 | 819200 | 50 | 256 | 52 | 1 | 0.4 | 0.001 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.049 | 0.272 |  | 0.416 | 1.7 | 1.27 |  | 2.808
case02 | 819200 | 50 | 256 | 52 | 1 | 0.4 | 0.0004 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.049 | 1.369 |  | 0.977 | 1.7 | -0.758 |  | 3.109
case02 | 819200 | 50 | 256 | 52 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.049 | 4.6 |  | 1.919 | 1.7 | -8.007 |  | 3.612 | 摸底
case02 | 819200 | 50 | 256 | 55 | 1 | 0.4 | 0.05 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.026 | 0.018 |  | 0.059 | -2.287 | -2.98 |  | -6.126
case02 | 819200 | 50 | 256 | 55 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.026 | 0.363 |  | 0.578 | -2.287 | -2.508 |  | -5.328
case02 | 819200 | 50 | 256 | 55 | 1 | 0.4 | 0.001 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.026 | 0.817 |  | 1.118 | -2.287 | -2.202 |  | -4.497
case02 | 819200 | 50 | 256 | 55 | 1 | 0.4 | 0.0004 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.026 | 2.727 |  | 2.741 | -2.287 | -2.421 |  | -2
case02 | 819200 | 50 | 256 | 55 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.026 | 7.78 |  | 5.449 | -2.287 | -6.594 |  | 2.158 | 摸底
case02 | 983040 | 60 | 128 | 55 | 1 | 0.4 | 0.05 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.074 | -0.1 |  | -0.195 | -1.613 | -1.639 |  | -3.268
case02 | 983040 | 60 | 128 | 55 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.074 | -0.797 |  | -1.774 | -1.613 | -3.133 |  | -4.776
case02 | 983040 | 60 | 128 | 55 | 1 | 0.4 | 0.001 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.074 | -2.047 |  | -3.442 | -1.613 | -6.397 |  | -6.348
case02 | 983040 | 60 | 128 | 55 | 1 | 0.4 | 0.0004 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.074 | -8.037 |  | -8.444 | -1.613 | -26.849 |  | -11.057
case02 | 983040 | 60 | 128 | 55 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.074 | -30.083 |  | -16.678 | -1.613 | -99.597 |  | -18.91 | 摸底
case02 | 983040 | 60 | 128 | 58 | 1 | 0.4 | 0.05 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.36 | 0.228 |  | -0.133 | -1.263 | -0.442 |  | -1.707
case02 | 983040 | 60 | 128 | 58 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.36 | -1.253 |  | -2.429 | -1.263 | -2.367 |  | -3.846
case02 | 983040 | 60 | 128 | 58 | 1 | 0.4 | 0.001 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.36 | -3.208 |  | -4.848 | -1.263 | -6.083 |  | -6.074
case02 | 983040 | 60 | 128 | 58 | 1 | 0.4 | 0.0004 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.36 | -12.462 |  | -11.994 | -1.263 | -27.95 |  | -12.754
case02 | 983040 | 60 | 128 | 58 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.36 | -32.074 |  | -24.002 | -1.263 | -103.447 |  | -23.89 | 摸底
case02 | 983040 | 60 | 128 | 62 | 1 | 0.4 | 0.05 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.423 | 0.327 |  | -0.116 | 0.164 | 0.896 |  | 1.054
case02 | 983040 | 60 | 128 | 62 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.423 | -0.56 |  | -2.191 | 0.164 | -1.171 |  | -1.081
case02 | 983040 | 60 | 128 | 62 | 1 | 0.4 | 0.001 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.423 | -1.883 |  | -4.365 | 0.164 | -5.038 |  | -3.307
case02 | 983040 | 60 | 128 | 62 | 1 | 0.4 | 0.0004 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.423 | -8.008 |  | -10.817 | 0.164 | -27.381 |  | -9.982
case02 | 983040 | 60 | 128 | 62 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.423 | -32.692 |  | -21.752 | 0.164 | -103.767 |  | -21.083 | 摸底
case02 | 983040 | 60 | 128 | 65 | 1 | 0.4 | 0.05 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.043 | -0.075 |  | -0.147 | -0.338 | -0.353 |  | -0.701
case02 | 983040 | 60 | 128 | 65 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.043 | -0.892 |  | -1.85 | -0.338 | -1.484 |  | -1.569
case02 | 983040 | 60 | 128 | 65 | 1 | 0.4 | 0.001 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.043 | -2.13 |  | -3.644 | -0.338 | -4.375 |  | -2.472
case02 | 983040 | 60 | 128 | 65 | 1 | 0.4 | 0.0004 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.043 | -7.473 |  | -8.957 | -0.338 | -23.738 |  | -5.186
case02 | 983040 | 60 | 128 | 65 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.043 | -30.627 |  | -17.864 | -0.338 | -94.557 |  | -9.668 | 摸底
case02 | 983040 | 60 | 256 | 55 | 1 | 0.4 | 0.05 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.043 | 0.014 |  | -0.104 | -2.854 | -2.799 |  | -6.147
case02 | 983040 | 60 | 256 | 55 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.043 | 1.541 |  | 1.475 | -2.854 | -1.448 |  | -4.542
case02 | 983040 | 60 | 256 | 55 | 1 | 0.4 | 0.001 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.043 | 3.212 |  | 3.122 | -2.854 | -0.231 |  | -2.87
case02 | 983040 | 60 | 256 | 55 | 1 | 0.4 | 0.0004 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.043 | 8.702 |  | 8.055 | -2.854 | 2.252 |  | 2.144
case02 | 983040 | 60 | 256 | 55 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.043 | 19.471 |  | 16.287 | -2.854 | 2.54 |  | 10.507 | 摸底
case02 | 983040 | 60 | 256 | 58 | 1 | 0.4 | 0.05 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.013 | -0.023 |  | -0.059 | -1.333 | -1.311 |  | -3.246
case02 | 983040 | 60 | 256 | 58 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.013 | -0.525 |  | -0.481 | -1.333 | -1.941 |  | -3.666
case02 | 983040 | 60 | 256 | 58 | 1 | 0.4 | 0.001 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.013 | -0.976 |  | -0.921 | -1.333 | -2.788 |  | -4.104
case02 | 983040 | 60 | 256 | 58 | 1 | 0.4 | 0.0004 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.013 | -1.761 |  | -2.241 | -1.333 | -6.509 |  | -5.415
case02 | 983040 | 60 | 256 | 58 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.013 | -1.604 |  | -4.437 | -1.333 | -16.692 |  | -7.6 | 摸底
case02 | 983040 | 60 | 256 | 62 | 1 | 0.4 | 0.05 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.099 | 0.145 |  | 0.054 | 1.441 | 1.682 |  | 2.389
case02 | 983040 | 60 | 256 | 62 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.099 | 0.4 |  | 0.477 | 1.441 | 2.003 |  | 3.008
case02 | 983040 | 60 | 256 | 62 | 1 | 0.4 | 0.001 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.099 | 0.741 |  | 0.919 | 1.441 | 2.15 |  | 3.654
case02 | 983040 | 60 | 256 | 62 | 1 | 0.4 | 0.0004 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.099 | 2.246 |  | 2.242 | 1.441 | 1.426 |  | 5.591
case02 | 983040 | 60 | 256 | 62 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | -0.099 | 6.328 |  | 4.441 | 1.441 | -3.681 |  | 8.807 | 摸底
case02 | 983040 | 60 | 256 | 65 | 1 | 0.4 | 0.05 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.019 | -0.013 |  | 0.013 | -0.043 | -0.084 |  | -0.969
case02 | 983040 | 60 | 256 | 65 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.019 | -0.675 |  | -0.663 | -0.043 | -1.117 |  | -1.847
case02 | 983040 | 60 | 256 | 65 | 1 | 0.4 | 0.001 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.019 | -1.266 |  | -1.366 | -0.043 | -2.38 |  | -2.76
case02 | 983040 | 60 | 256 | 65 | 1 | 0.4 | 0.0004 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.019 | -2.505 |  | -3.47 | -0.043 | -7.325 |  | -5.495
case02 | 983040 | 60 | 256 | 65 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.019 | -2.866 |  | -6.993 | -0.043 | -19.514 |  | -10.071 | 摸底

## case03 test report
dynamic_test | ADC_samp_freq | fund_wave_freq | ADC_OSR | scan_freq | PGA | fund_wave_amp_vol | fund_wave_amp_current | power_factor | vol_dc_bias | current_dc_bias | harmonic_wave_oder | amp_harmonic_vol | amp_harmonic_current | full_RMSU_err (RMSU电压有效值误差(‰)) | full_RMSI_err (RMSI电流有效值误差(‰)) | full_P_err (P有功功率误差(‰)) | full_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | 备注
case03 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 | 0.045 | 0.339 | 0.359 |  | 0.074 | 0.337 | 0.426
case03 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 2 | 0 | 0 | 0.045 | 0.681 | 0.699 |  | 0.074 | 0.674 | 0.725
case03 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.045 | 4.394 | 4.262 |  | 0.074 | 4.225 | 4.239 |  | 摸底
case03 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.045 | 49.401 | 42.466 |  | 0.074 | 120.661 | 42.242 |  | 摸底
case03 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 | 0.01 | 0.044 | 0.107 |  | 0.256 | 0.271 | 0.596
case03 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 2 | 0 | 0 | 0.01 | 0.089 | 0.181 |  | 0.256 | 0.286 | 0.668
case03 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.01 | 0.557 | 0.96 |  | 0.256 | 0.676 | 1.415 |  | 摸底
case03 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.01 | 8.22 | 9.299 |  | 0.256 | 14.125 | 9.42 |  | 摸底
case03 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 | 0.044 | 0.339 | 0.359 |  | 0.074 | 0.336 | 0.426
case03 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 2 | 0 | 0 | 0.044 | 0.678 | 0.698 |  | 0.074 | 0.674 | 0.724
case03 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.044 | 4.375 | 4.261 |  | 0.074 | 4.224 | 4.237 |  | 摸底
case03 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.044 | 44.178 | 42.465 |  | 0.074 | 120.655 | 42.238 |  | 摸底
case03 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 | 0.01 | 0.045 | 0.108 |  | 0.256 | 0.271 | 0.596
case03 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 2 | 0 | 0 | 0.01 | 0.089 | 0.182 |  | 0.256 | 0.285 | 0.667
case03 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.01 | 0.554 | 0.96 |  | 0.256 | 0.676 | 1.415 |  | 摸底
case03 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.01 | 8.227 | 9.305 |  | 0.256 | 14.126 | 9.418 |  | 摸底
case03 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.025 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.045 | 0.158 | 0.68 | 0.271 | 0.074 | 0.207 | 0.722 | 0.336
case03 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0125 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.045 | 0.314 | 1.359 | 0.563 | 0.074 | 0.415 | 1.35 | 0.602
case03 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.045 | 1.961 | 8.483 | 3.565 | 0.074 | 2.6 | 8.434 | 3.565 | 摸底
case03 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.045 | 28.405 | 84.74 | 35.812 | 0.074 | 102.768 | 84.232 | 35.831 | 摸底
case03 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.025 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.01 | 0.087 | 0.176 | 0.101 | 0.256 | 0.306 | 0.673 | 0.232
case03 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0125 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.01 | 0.173 | 0.324 | 0.203 | 0.256 | 0.355 | 0.816 | 0.299
case03 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.01 | 1.078 | 1.887 | 1.272 | 0.256 | 1.483 | 2.311 | 1.862 | 摸底
case03 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.01 | 11.554 | 18.712 | 12.632 | 0.256 | 17.801 | 18.361 | 18.555 | 摸底
case03 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.025 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.044 | 0.158 | 0.68 | 0.271 | 0.074 | 0.207 | 0.721 | 0.336
case03 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0125 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.044 | 0.314 | 1.358 | 0.562 | 0.074 | 0.415 | 1.349 | 0.602
case03 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.044 | 2.009 | 8.481 | 3.579 | 0.074 | 2.599 | 8.433 | 3.564 | 摸底
case03 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.044 | 28.584 | 84.74 | 35.75 | 0.074 | 102.766 | 84.235 | 35.834 | 摸底
case03 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.025 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.01 | 0.087 | 0.177 | 0.101 | 0.256 | 0.306 | 0.673 | 0.232
case03 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0125 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.01 | 0.173 | 0.324 | 0.203 | 0.256 | 0.355 | 0.815 | 0.299
case03 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.01 | 1.084 | 1.888 | 1.272 | 0.256 | 1.483 | 2.311 | 1.862 | 摸底
case03 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.01 | 11.51 | 18.719 | 12.634 | 0.256 | 17.801 | 18.36 | 18.554 | 摸底
case03 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.025 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.045 | 0.226 |  | 0.245 | 0.074 | 0.252 |  | 0.306
case03 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0125 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.045 | 0.454 |  | 0.495 | 0.074 | 0.504 |  | 0.536
case03 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.045 | 2.8 |  | 3.081 | 0.074 | 3.155 |  | 3.093 | 摸底
case03 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.045 | 29.583 |  | 31.156 | 0.074 | 111.473 |  | 31.001 | 摸底
case03 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.025 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.01 | 0.099 |  | 0.09 | 0.256 | 0.304 |  | 0.226
case03 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0125 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.01 | 0.198 |  | 0.179 | 0.256 | 0.352 |  | 0.284
case03 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.01 | 1.232 |  | 1.108 | 0.256 | 1.324 |  | 1.62 | 摸底
case03 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.01 | 12.202 |  | 11.074 | 0.256 | 16.428 |  | 16.178 | 摸底
case03 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.025 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.044 | 0.225 |  | 0.243 | 0.074 | 0.252 |  | 0.306
case03 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0125 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.044 | 0.453 |  | 0.494 | 0.074 | 0.504 |  | 0.536
case03 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.044 | 2.754 |  | 3.108 | 0.074 | 3.155 |  | 3.093 | 摸底
case03 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.044 | 30.376 |  | 31.239 | 0.074 | 111.472 |  | 31.001 | 摸底
case03 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.025 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.01 | 0.099 |  | 0.09 | 0.256 | 0.305 |  | 0.226
case03 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0125 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.01 | 0.198 |  | 0.179 | 0.256 | 0.352 |  | 0.284
case03 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.01 | 1.23 |  | 1.109 | 0.256 | 1.324 |  | 1.62 | 摸底
case03 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0002 | 6.1232e-17 | 0 | 0 | 2 | 0 | 0 | 0.01 | 12.606 |  | 11.072 | 0.256 | 16.427 |  | 16.176 | 摸底

## case04 test report
dynamic_test | ADC_samp_freq | fund_wave_freq | ADC_OSR | scan_freq | PGA | fund_wave_amp_vol | fund_wave_amp_current | power_factor | vol_dc_bias | current_dc_bias | harmonic_wave_oder | amp_harmonic_vol | amp_harmonic_current | full_RMSU_err (RMSU电压有效值误差(‰)) | full_RMSI_err (RMSI电流有效值误差(‰)) | full_P_err (P有功功率误差(‰)) | full_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | 备注
case04 | 819200 | 50 | 128 | 50 | 1 | 0.24 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.024 | -0.031 | -0.052 |  | -0.057 | -0.06 | -0.113 |  | 摸底
case04 | 819200 | 50 | 128 | 50 | 1 | 0.24 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.024 | 0.01 | -0.08 |  | -0.057 | -0.045 | -0.167 |  | 摸底
case04 | 819200 | 50 | 128 | 50 | 1 | 0.24 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.024 | 0.99 | -0.776 |  | -0.057 | -0.491 | -1.515 |  | 摸底
case04 | 819200 | 50 | 128 | 50 | 1 | 0.24 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.024 | -13.722 | -7.283 |  | -0.057 | -85.915 | -14.143 |  | 摸底
case04 | 819200 | 50 | 128 | 50 | 1 | 0.32 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.027 | -0.031 | -0.054 |  | -0.058 | -0.06 | -0.113
case04 | 819200 | 50 | 128 | 50 | 1 | 0.32 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.027 | 0.01 | -0.064 |  | -0.058 | -0.045 | -0.149
case04 | 819200 | 50 | 128 | 50 | 1 | 0.32 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.027 | 0.983 | -0.32 |  | -0.058 | -0.491 | -1.056
case04 | 819200 | 50 | 128 | 50 | 1 | 0.32 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.027 | -14.191 | -2.715 |  | -0.058 | -85.915 | -9.548 |  | 摸底
case04 | 819200 | 50 | 128 | 50 | 1 | 0.48 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.029 | -0.031 | -0.055 |  | -0.059 | -0.06 | -0.113
case04 | 819200 | 50 | 128 | 50 | 1 | 0.48 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.029 | 0.01 | -0.048 |  | -0.059 | -0.045 | -0.132
case04 | 819200 | 50 | 128 | 50 | 1 | 0.48 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.029 | 0.986 | 0.134 |  | -0.059 | -0.491 | -0.598
case04 | 819200 | 50 | 128 | 50 | 1 | 0.48 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.029 | -14.263 | 1.857 |  | -0.059 | -85.915 | -4.953 |  | 摸底
case04 | 819200 | 50 | 256 | 50 | 1 | 0.24 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.003 | 0.008 | 0.027 |  | 0.251 | 0.254 | 0.522 |  | 摸底
case04 | 819200 | 50 | 256 | 50 | 1 | 0.24 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.003 | -0.018 | 0.044 |  | 0.251 | 0.233 | 0.544 |  | 摸底
case04 | 819200 | 50 | 256 | 50 | 1 | 0.24 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.003 | -0.615 | 0.489 |  | 0.251 | -0.395 | 1.107 |  | 摸底
case04 | 819200 | 50 | 256 | 50 | 1 | 0.24 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.003 | -2.213 | 4.666 |  | 0.251 | -15.137 | 6.396 |  | 摸底
case04 | 819200 | 50 | 256 | 50 | 1 | 0.32 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.008 | 0.028 |  | 0.252 | 0.254 | 0.522
case04 | 819200 | 50 | 256 | 50 | 1 | 0.32 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.018 | 0.034 |  | 0.252 | 0.233 | 0.533
case04 | 819200 | 50 | 256 | 50 | 1 | 0.32 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.617 | 0.205 |  | 0.252 | -0.395 | 0.823
case04 | 819200 | 50 | 256 | 50 | 1 | 0.32 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -2.214 | 1.812 |  | 0.252 | -15.137 | 3.547 |  | 摸底
case04 | 819200 | 50 | 256 | 50 | 1 | 0.48 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.006 | 0.008 | 0.029 |  | 0.253 | 0.254 | 0.523
case04 | 819200 | 50 | 256 | 50 | 1 | 0.48 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.006 | -0.018 | 0.024 |  | 0.253 | 0.233 | 0.523
case04 | 819200 | 50 | 256 | 50 | 1 | 0.48 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.006 | -0.616 | -0.079 |  | 0.253 | -0.395 | 0.539
case04 | 819200 | 50 | 256 | 50 | 1 | 0.48 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.006 | -2.211 | -1.041 |  | 0.253 | -15.137 | 0.699 |  | 摸底
case04 | 983040 | 60 | 128 | 60 | 1 | 0.24 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.024 | -0.031 | -0.052 |  | -0.058 | -0.06 | -0.115 |  | 摸底
case04 | 983040 | 60 | 128 | 60 | 1 | 0.24 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.024 | 0.011 | -0.079 |  | -0.058 | -0.046 | -0.169 |  | 摸底
case04 | 983040 | 60 | 128 | 60 | 1 | 0.24 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.024 | 0.945 | -0.775 |  | -0.058 | -0.492 | -1.516 |  | 摸底
case04 | 983040 | 60 | 128 | 60 | 1 | 0.24 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.024 | -9.231 | -7.29 |  | -0.058 | -85.926 | -14.154 |  | 摸底
case04 | 983040 | 60 | 128 | 60 | 1 | 0.32 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.026 | -0.031 | -0.053 |  | -0.059 | -0.06 | -0.115
case04 | 983040 | 60 | 128 | 60 | 1 | 0.32 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.026 | 0.011 | -0.063 |  | -0.059 | -0.046 | -0.151
case04 | 983040 | 60 | 128 | 60 | 1 | 0.32 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.026 | 0.966 | -0.32 |  | -0.059 | -0.492 | -1.058
case04 | 983040 | 60 | 128 | 60 | 1 | 0.32 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.026 | -8.816 | -2.719 |  | -0.059 | -85.926 | -9.56 |  | 摸底
case04 | 983040 | 60 | 128 | 60 | 1 | 0.48 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.029 | -0.031 | -0.054 |  | -0.06 | -0.06 | -0.115
case04 | 983040 | 60 | 128 | 60 | 1 | 0.48 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.029 | 0.011 | -0.047 |  | -0.06 | -0.046 | -0.134
case04 | 983040 | 60 | 128 | 60 | 1 | 0.48 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.029 | 0.946 | 0.135 |  | -0.06 | -0.492 | -0.599
case04 | 983040 | 60 | 128 | 60 | 1 | 0.48 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.029 | -8.732 | 1.851 |  | -0.06 | -85.926 | -4.964 |  | 摸底
case04 | 983040 | 60 | 256 | 60 | 1 | 0.24 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.003 | 0.008 | 0.027 |  | 0.251 | 0.254 | 0.521 |  | 摸底
case04 | 983040 | 60 | 256 | 60 | 1 | 0.24 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.003 | -0.018 | 0.044 |  | 0.251 | 0.233 | 0.543 |  | 摸底
case04 | 983040 | 60 | 256 | 60 | 1 | 0.24 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.003 | -0.618 | 0.489 |  | 0.251 | -0.396 | 1.107 |  | 摸底
case04 | 983040 | 60 | 256 | 60 | 1 | 0.24 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.003 | -2.268 | 4.662 |  | 0.251 | -15.143 | 6.391 |  | 摸底
case04 | 983040 | 60 | 256 | 60 | 1 | 0.32 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.008 | 0.028 |  | 0.252 | 0.254 | 0.522
case04 | 983040 | 60 | 256 | 60 | 1 | 0.32 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.018 | 0.034 |  | 0.252 | 0.233 | 0.533
case04 | 983040 | 60 | 256 | 60 | 1 | 0.32 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.616 | 0.205 |  | 0.252 | -0.396 | 0.823
case04 | 983040 | 60 | 256 | 60 | 1 | 0.32 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -2.26 | 1.808 |  | 0.252 | -15.142 | 3.542 |  | 摸底
case04 | 983040 | 60 | 256 | 60 | 1 | 0.48 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.006 | 0.008 | 0.029 |  | 0.253 | 0.254 | 0.523
case04 | 983040 | 60 | 256 | 60 | 1 | 0.48 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.006 | -0.018 | 0.025 |  | 0.253 | 0.233 | 0.523
case04 | 983040 | 60 | 256 | 60 | 1 | 0.48 | 0.002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.006 | -0.618 | -0.079 |  | 0.253 | -0.396 | 0.539
case04 | 983040 | 60 | 256 | 60 | 1 | 0.48 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.006 | -2.291 | -1.047 |  | 0.253 | -15.143 | 0.693 |  | 摸底
case04 | 819200 | 50 | 128 | 50 | 1 | 0.24 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.024 | -0.03 | -0.036 | -0.025 | -0.057 | -0.058 | -0.102 | -0.086 | 摸底
case04 | 819200 | 50 | 128 | 50 | 1 | 0.24 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.024 | 0.033 | -0.091 | -0.057 | -0.057 | -0.019 | -0.209 | -0.117 | 摸底
case04 | 819200 | 50 | 128 | 50 | 1 | 0.24 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.024 | 1.579 | -1.489 | -0.863 | -0.057 | 0.148 | -2.911 | -0.902 | 摸底
case04 | 819200 | 50 | 128 | 50 | 1 | 0.24 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.024 | -4.157 | -14.622 | -8.344 | -0.057 | -78.964 | -28.27 | -8.236 | 摸底
case04 | 819200 | 50 | 128 | 50 | 1 | 0.32 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.027 | -0.03 | -0.041 | -0.025 | -0.058 | -0.058 | -0.104 | -0.084
case04 | 819200 | 50 | 128 | 50 | 1 | 0.32 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.027 | 0.033 | -0.062 | -0.037 | -0.058 | -0.019 | -0.176 | -0.095
case04 | 819200 | 50 | 128 | 50 | 1 | 0.32 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.027 | 1.605 | -0.58 | -0.34 | -0.058 | 0.148 | -1.996 | -0.381
case04 | 819200 | 50 | 128 | 50 | 1 | 0.32 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.027 | -4.073 | -5.484 | -3.142 | -0.058 | -78.964 | -19.082 | -3.032 | 摸底
case04 | 819200 | 50 | 128 | 50 | 1 | 0.48 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.029 | -0.03 | -0.046 | -0.024 | -0.059 | -0.058 | -0.107 | -0.082
case04 | 819200 | 50 | 128 | 50 | 1 | 0.48 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.029 | 0.033 | -0.032 | -0.016 | -0.059 | -0.019 | -0.144 | -0.074
case04 | 819200 | 50 | 128 | 50 | 1 | 0.48 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.029 | 1.582 | 0.327 | 0.176 | -0.059 | 0.148 | -1.082 | 0.14
case04 | 819200 | 50 | 128 | 50 | 1 | 0.48 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.029 | -4.102 | 3.655 | 2.071 | -0.059 | -78.964 | -9.894 | 2.172 | 摸底
case04 | 819200 | 50 | 256 | 50 | 1 | 0.24 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.003 | 0.008 | 0.026 | 0.002 | 0.251 | 0.255 | 0.521 | -0.168 | 摸底
case04 | 819200 | 50 | 256 | 50 | 1 | 0.24 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.003 | -0.009 | 0.061 | 0.006 | 0.251 | 0.239 | 0.565 | -0.162 | 摸底
case04 | 819200 | 50 | 256 | 50 | 1 | 0.24 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.003 | -0.374 | 0.946 | 0.104 | 0.251 | -0.233 | 1.686 | -0.017 | 摸底
case04 | 819200 | 50 | 256 | 50 | 1 | 0.24 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.003 | -0.013 | 9.136 | 1.079 | 0.251 | -13.471 | 12.082 | 1.409 | 摸底
case04 | 819200 | 50 | 256 | 50 | 1 | 0.32 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.008 | 0.027 | 0.003 | 0.252 | 0.255 | 0.521 | -0.167
case04 | 819200 | 50 | 256 | 50 | 1 | 0.32 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.009 | 0.04 | 0.004 | 0.252 | 0.239 | 0.544 | -0.164
case04 | 819200 | 50 | 256 | 50 | 1 | 0.32 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.372 | 0.376 | 0.037 | 0.252 | -0.234 | 1.118 | -0.083
case04 | 819200 | 50 | 256 | 50 | 1 | 0.32 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.075 | 3.428 | 0.404 | 0.252 | -13.472 | 6.384 | 0.734 | 摸底
case04 | 819200 | 50 | 256 | 50 | 1 | 0.48 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.006 | 0.008 | 0.027 | 0.004 | 0.253 | 0.255 | 0.522 | -0.166
case04 | 819200 | 50 | 256 | 50 | 1 | 0.48 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.006 | -0.009 | 0.019 | 0.003 | 0.253 | 0.239 | 0.523 | -0.166
case04 | 819200 | 50 | 256 | 50 | 1 | 0.48 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.006 | -0.374 | -0.193 | -0.029 | 0.253 | -0.234 | 0.55 | -0.15
case04 | 819200 | 50 | 256 | 50 | 1 | 0.48 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.006 | 0.002 | -2.283 | -0.269 | 0.253 | -13.472 | 0.685 | 0.061 | 摸底
case04 | 983040 | 60 | 128 | 60 | 1 | 0.24 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.024 | -0.03 | -0.034 | -0.025 | -0.058 | -0.059 | -0.104 | -0.087 | 摸底
case04 | 983040 | 60 | 128 | 60 | 1 | 0.24 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.024 | 0.033 | -0.09 | -0.057 | -0.058 | -0.02 | -0.211 | -0.119 | 摸底
case04 | 983040 | 60 | 128 | 60 | 1 | 0.24 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.024 | 1.533 | -1.487 | -0.859 | -0.058 | 0.147 | -2.914 | -0.904 | 摸底
case04 | 983040 | 60 | 128 | 60 | 1 | 0.24 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.024 | -5.467 | -14.608 | -8.389 | -0.058 | -78.971 | -28.264 | -8.248 | 摸底
case04 | 983040 | 60 | 128 | 60 | 1 | 0.32 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.026 | -0.03 | -0.039 | -0.025 | -0.059 | -0.059 | -0.106 | -0.085
case04 | 983040 | 60 | 128 | 60 | 1 | 0.32 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.026 | 0.034 | -0.06 | -0.037 | -0.059 | -0.02 | -0.178 | -0.097
case04 | 983040 | 60 | 128 | 60 | 1 | 0.32 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.026 | 1.531 | -0.579 | -0.338 | -0.059 | 0.147 | -1.999 | -0.382
case04 | 983040 | 60 | 128 | 60 | 1 | 0.32 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.026 | -5.242 | -5.47 | -3.205 | -0.059 | -78.971 | -19.075 | -3.044 | 摸底
case04 | 983040 | 60 | 128 | 60 | 1 | 0.48 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.029 | -0.03 | -0.044 | -0.024 | -0.06 | -0.059 | -0.109 | -0.084
case04 | 983040 | 60 | 128 | 60 | 1 | 0.48 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.029 | 0.033 | -0.03 | -0.017 | -0.06 | -0.02 | -0.146 | -0.075
case04 | 983040 | 60 | 128 | 60 | 1 | 0.48 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.029 | 1.54 | 0.328 | 0.185 | -0.06 | 0.147 | -1.084 | 0.139
case04 | 983040 | 60 | 128 | 60 | 1 | 0.48 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | -0.029 | -5.44 | 3.669 | 2.01 | -0.06 | -78.971 | -9.887 | 2.16 | 摸底
case04 | 983040 | 60 | 256 | 60 | 1 | 0.24 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.003 | 0.008 | 0.027 | 0.002 | 0.251 | 0.255 | 0.521 | -0.168 | 摸底
case04 | 983040 | 60 | 256 | 60 | 1 | 0.24 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.003 | -0.008 | 0.062 | 0.006 | 0.251 | 0.239 | 0.565 | -0.163 | 摸底
case04 | 983040 | 60 | 256 | 60 | 1 | 0.24 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.003 | -0.372 | 0.946 | 0.105 | 0.251 | -0.233 | 1.686 | -0.016 | 摸底
case04 | 983040 | 60 | 256 | 60 | 1 | 0.24 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.003 | 0.243 | 9.135 | 1.08 | 0.251 | -13.471 | 12.079 | 1.411 | 摸底
case04 | 983040 | 60 | 256 | 60 | 1 | 0.32 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.008 | 0.027 | 0.003 | 0.252 | 0.255 | 0.521 | -0.167
case04 | 983040 | 60 | 256 | 60 | 1 | 0.32 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.008 | 0.04 | 0.004 | 0.252 | 0.239 | 0.544 | -0.164
case04 | 983040 | 60 | 256 | 60 | 1 | 0.32 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.372 | 0.377 | 0.038 | 0.252 | -0.233 | 1.119 | -0.083
case04 | 983040 | 60 | 256 | 60 | 1 | 0.32 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.258 | 3.425 | 0.406 | 0.252 | -13.471 | 6.38 | 0.736 | 摸底
case04 | 983040 | 60 | 256 | 60 | 1 | 0.48 | 1 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.006 | 0.008 | 0.028 | 0.004 | 0.253 | 0.255 | 0.522 | -0.167
case04 | 983040 | 60 | 256 | 60 | 1 | 0.48 | 0.05 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.006 | -0.008 | 0.019 | 0.003 | 0.253 | 0.239 | 0.523 | -0.166
case04 | 983040 | 60 | 256 | 60 | 1 | 0.48 | 0.002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.006 | -0.372 | -0.192 | -0.028 | 0.253 | -0.233 | 0.551 | -0.149
case04 | 983040 | 60 | 256 | 60 | 1 | 0.48 | 0.0002 | 0.5 | 0 | 0 | 2 | 0 | 0 | 0.006 | 0.266 | -2.283 | -0.268 | 0.253 | -13.472 | 0.682 | 0.062 | 摸底

## case05 test report
dynamic_test | ADC_samp_freq | fund_wave_freq | ADC_OSR | scan_freq | PGA | fund_wave_amp_vol | fund_wave_amp_current | power_factor | vol_dc_bias | current_dc_bias | harmonic_wave_oder | amp_harmonic_vol | amp_harmonic_current | full_RMSU_err (RMSU电压有效值误差(‰)) | full_RMSI_err (RMSI电流有效值误差(‰)) | full_P_err (P有功功率误差(‰)) | full_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | 备注
case05 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.031 | -0.054 |  | -0.058 | -0.06 | -0.113
case05 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.5 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.029 | -0.054 |  | -0.058 | -0.059 | -0.115
case05 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.012 | -0.054 |  | -0.058 | -0.052 | -0.125
case05 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.01 | -0.054 |  | -0.058 | -0.045 | -0.139
case05 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.005 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.411 | -0.052 |  | -0.058 | -0.025 | -0.38
case05 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0025 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.827 | -0.049 |  | -0.058 | -0.267 | -0.648
case05 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 1.51 | -0.042 |  | -0.058 | -2.656 | -1.452
case05 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 2.216 | -0.025 |  | -0.058 | -12.235 | -2.789
case05 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -13.637 | 0.028 |  | -0.058 | -85.916 | -6.792
case05 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.008 | 0.028 |  | 0.253 | 0.254 | 0.523
case05 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.5 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.006 | 0.028 |  | 0.253 | 0.253 | 0.523
case05 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.005 | 0.028 |  | 0.253 | 0.244 | 0.525
case05 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case05 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.005 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.255 | 0.031 |  | 0.253 | 0.018 | 0.574
case05 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0025 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.498 | 0.034 |  | 0.253 | -0.249 | 0.627
case05 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -1.16 | 0.042 |  | 0.253 | -1.241 | 0.784
case05 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -1.998 | 0.056 |  | 0.253 | -3.522 | 1.047
case05 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -2.21 | 0.101 |  | 0.253 | -15.137 | 1.839
case05 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.031 | -0.054 |  | -0.059 | -0.06 | -0.115
case05 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.5 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.029 | -0.054 |  | -0.059 | -0.06 | -0.116
case05 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.1 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -0.011 | -0.054 |  | -0.059 | -0.053 | -0.127
case05 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.011 | -0.054 |  | -0.059 | -0.046 | -0.141
case05 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.005 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.406 | -0.052 |  | -0.059 | -0.027 | -0.382
case05 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0025 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.803 | -0.049 |  | -0.059 | -0.269 | -0.65
case05 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 1.464 | -0.039 |  | -0.059 | -2.655 | -1.452
case05 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 2.054 | -0.027 |  | -0.059 | -12.236 | -2.791
case05 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | -8.783 | 0.023 |  | -0.059 | -85.926 | -6.802
case05 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.008 | 0.029 |  | 0.253 | 0.254 | 0.522
case05 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.5 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | 0.006 | 0.028 |  | 0.253 | 0.253 | 0.522
case05 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.1 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.005 | 0.028 |  | 0.253 | 0.244 | 0.524
case05 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case05 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.005 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.256 | 0.031 |  | 0.253 | 0.018 | 0.574
case05 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0025 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -0.499 | 0.033 |  | 0.253 | -0.25 | 0.627
case05 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -1.162 | 0.041 |  | 0.253 | -1.242 | 0.783
case05 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -1.992 | 0.054 |  | 0.253 | -3.525 | 1.045
case05 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0002 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -2.256 | 0.096 |  | 0.253 | -15.142 | 1.833

## case06 test report
dynamic_test | ADC_samp_freq | fund_wave_freq | ADC_OSR | scan_freq | PGA | fund_wave_amp_vol | fund_wave_amp_current | power_factor | vol_dc_bias | current_dc_bias | harmonic_wave_oder | amp_harmonic_vol | amp_harmonic_current | full_RMSU_err (RMSU电压有效值误差(‰)) | full_RMSI_err (RMSI电流有效值误差(‰)) | full_P_err (P有功功率误差(‰)) | full_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | 备注
case06 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | -6.9863 |  |  |  | -8.0977
case06 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | 1.9991 |  |  |  | 9.6743
case06 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | -6.987 |  |  |  | -8.097
case06 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | 1.9963 |  |  |  | 9.6722

## case10 test report
dynamic_test | ADC_samp_freq | fund_wave_freq | ADC_OSR | scan_freq | PGA | fund_wave_amp_vol | fund_wave_amp_current | power_factor | vol_dc_bias | current_dc_bias | harmonic_wave_oder | amp_harmonic_vol | amp_harmonic_current | full_RMSU_err (RMSU电压有效值误差(‰)) | full_RMSI_err (RMSI电流有效值误差(‰)) | full_P_err (P有功功率误差(‰)) | full_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | 备注
case10 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 13 | 0 | 0 | -0.027 | 0.138 | -0.079 |  | -0.058 | -0.029 | -0.166
case10 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 13 | 0 | 0 | -0.027 | 0.189 | -0.075 |  | -0.058 | -0.011 | -0.22
case10 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 13 | 0 | 0 | -0.027 | 0.946 | 0.024 |  | -0.058 | -2.652 | -1.451 |  | 摸底
case10 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 13 | 0 | 0 | 0.007 | 0.119 | -0.006 |  | 0.253 | 0.212 | 0.532
case10 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 13 | 0 | 0 | 0.007 | 0.085 | -0.008 |  | 0.253 | 0.165 | 0.543
case10 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 13 | 0 | 0 | 0.007 | -0.524 | -0.051 |  | 0.253 | -1.241 | 0.783 |  | 摸底
case10 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 13 | 0 | 0 | -0.026 | 0.138 | -0.078 |  | -0.059 | -0.03 | -0.168
case10 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 13 | 0 | 0 | -0.026 | 0.189 | -0.074 |  | -0.059 | -0.012 | -0.222
case10 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 13 | 0 | 0 | -0.026 | 0.971 | 0.023 |  | -0.059 | -2.655 | -1.454 |  | 摸底
case10 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 13 | 0 | 0 | 0.007 | 0.119 | -0.006 |  | 0.253 | 0.211 | 0.532
case10 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 13 | 0 | 0 | 0.007 | 0.086 | -0.008 |  | 0.253 | 0.165 | 0.542
case10 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 13 | 0 | 0 | 0.007 | -0.532 | -0.048 |  | 0.253 | -1.241 | 0.783 |  | 摸底

## case11 test report
dynamic_test | ADC_samp_freq | fund_wave_freq | ADC_OSR | scan_freq | PGA | fund_wave_amp_vol | fund_wave_amp_current | power_factor | vol_dc_bias | current_dc_bias | harmonic_wave_oder | amp_harmonic_vol | amp_harmonic_current | full_RMSU_err (RMSU电压有效值误差(‰)) | full_RMSI_err (RMSI电流有效值误差(‰)) | full_P_err (P有功功率误差(‰)) | full_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | 备注
case11 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 13 | 0 | 0 | -0.025 | 0.17 | -0.095 |  | -0.058 | -0.029 | -0.166
case11 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 13 | 0 | 0 | -0.025 | 0.268 | -0.098 |  | -0.058 | -0.011 | -0.22
case11 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 13 | 0 | 0 | -0.025 | 1.793 | -0.155 |  | -0.058 | -2.651 | -1.45 |  | 摸底
case11 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 13 | 0 | 0 | 0.009 | 0.105 | -0.018 |  | 0.253 | 0.211 | 0.532
case11 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 13 | 0 | 0 | 0.009 | 0.072 | -0.02 |  | 0.253 | 0.165 | 0.542
case11 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 13 | 0 | 0 | 0.009 | -0.587 | -0.049 |  | 0.253 | -1.241 | 0.783 |  | 摸底
case11 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 13 | 0 | 0 | -0.025 | 0.17 | -0.095 |  | -0.059 | -0.03 | -0.168
case11 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 13 | 0 | 0 | -0.025 | 0.27 | -0.097 |  | -0.059 | -0.012 | -0.222
case11 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 13 | 0 | 0 | -0.025 | 1.836 | -0.155 |  | -0.059 | -2.653 | -1.453 |  | 摸底
case11 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 13 | 0 | 0 | 0.009 | 0.106 | -0.018 |  | 0.253 | 0.211 | 0.532
case11 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 13 | 0 | 0 | 0.009 | 0.071 | -0.02 |  | 0.253 | 0.165 | 0.542
case11 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 13 | 0 | 0 | 0.009 | -0.582 | -0.049 |  | 0.253 | -1.242 | 0.782 |  | 摸底

## case12 test report
dynamic_test | ADC_samp_freq | fund_wave_freq | ADC_OSR | scan_freq | PGA | fund_wave_amp_vol | fund_wave_amp_current | power_factor | vol_dc_bias | current_dc_bias | harmonic_wave_oder | amp_harmonic_vol | amp_harmonic_current | full_RMSU_err (RMSU电压有效值误差(‰)) | full_RMSI_err (RMSI电流有效值误差(‰)) | full_P_err (P有功功率误差(‰)) | full_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | 备注
case12 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | -2.5957 |  |  |  | -3.5355
case12 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | -3.9409 |  |  |  | -4.9308
case12 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | -34.8105 |  |  |  | -36.958 |  | 摸底
case12 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | 1.9974 |  |  |  | 9.677
case12 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | 3.4529 |  |  |  | 10.927
case12 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | 36.9301 |  |  |  | 39.6749 |  | 摸底
case12 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | -2.588 |  |  |  | -3.5076
case12 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | -3.9317 |  |  |  | -4.9017
case12 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | -34.7754 |  |  |  | -36.8837 |  | 摸底
case12 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | 1.9949 |  |  |  | 9.67
case12 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | 3.4483 |  |  |  | 10.9174
case12 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | 36.9346 |  |  |  | 39.6744 |  | 摸底

## case13 test report
dynamic_test | ADC_samp_freq | fund_wave_freq | ADC_OSR | scan_freq | PGA | fund_wave_amp_vol | fund_wave_amp_current | power_factor | vol_dc_bias | current_dc_bias | harmonic_wave_oder | amp_harmonic_vol | amp_harmonic_current | full_RMSU_err (RMSU电压有效值误差(‰)) | full_RMSI_err (RMSI电流有效值误差(‰)) | full_P_err (P有功功率误差(‰)) | full_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | 备注
case13 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.189 | -0.066 |  | -0.058 | -0.023 | -0.165
case13 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.269 | -0.066 |  | -0.058 | -0.006 | -0.219
case13 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 1.512 | -0.049 |  | -0.058 | -2.644 | -1.448 |  | 摸底
case13 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -1.37 | 0.014 |  | 0.253 | 0.214 | 0.532
case13 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -1.418 | 0.014 |  | 0.253 | 0.168 | 0.542
case13 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -2.572 | 0.029 |  | 0.253 | -1.234 | 0.787 |  | 摸底
case13 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.189 | -0.066 |  | -0.059 | -0.024 | -0.167
case13 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 0.269 | -0.065 |  | -0.059 | -0.006 | -0.221
case13 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | -0.028 | 1.525 | -0.049 |  | -0.059 | -2.646 | -1.451 |  | 摸底
case13 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -1.371 | 0.014 |  | 0.253 | 0.214 | 0.532
case13 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -1.417 | 0.014 |  | 0.253 | 0.168 | 0.542
case13 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 | 0.005 | -2.523 | 0.029 |  | 0.253 | -1.235 | 0.787 |  | 摸底

## case14 test report
dynamic_test | ADC_samp_freq | fund_wave_freq | ADC_OSR | scan_freq | PGA | fund_wave_amp_vol | fund_wave_amp_current | power_factor | vol_dc_bias | current_dc_bias | harmonic_wave_oder | amp_harmonic_vol | amp_harmonic_current | full_RMSU_err (RMSU电压有效值误差(‰)) | full_RMSI_err (RMSI电流有效值误差(‰)) | full_P_err (P有功功率误差(‰)) | full_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | 备注
case14 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | -2.2436 |  |  |  | -3.2005
case14 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | -3.5692 |  |  |  | -4.5773
case14 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | -34.4286 |  |  |  | -36.5782
case14 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | 1.9644 |  |  |  | 9.6405
case14 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | 3.4134 |  |  |  | 10.8831
case14 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | 36.9693 |  |  |  | 39.7247
case14 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | -2.2584 |  |  |  | -3.2134
case14 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | -3.585 |  |  |  | -4.5906
case14 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | -34.4741 |  |  |  | -36.624
case14 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.025 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | 1.9483 |  |  |  | 9.627
case14 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0125 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | 3.3962 |  |  |  | 10.8686
case14 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 2 | 0 | 0 |  |  | 36.9385 |  |  |  | 39.6981

## case15 test report
dynamic_test | ADC_samp_freq | fund_wave_freq | ADC_OSR | scan_freq | PGA | fund_wave_amp_vol | fund_wave_amp_current | power_factor | vol_dc_bias | current_dc_bias | harmonic_wave_oder | amp_harmonic_vol | amp_harmonic_current | full_RMSU_err (RMSU电压有效值误差(‰)) | full_RMSI_err (RMSI电流有效值误差(‰)) | full_P_err (P有功功率误差(‰)) | full_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | 备注
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 15 | 0.008 | 0 | -0.026 | 0.01 | -0.053 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 15 | 0 | 0.005 | -0.028 | 0.051 | -0.053 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 16 | 0.008 | 0 | -0.027 | 0.01 | -0.055 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 16 | 0 | 0.005 | -0.028 | 0.036 | -0.055 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 17 | 0.008 | 0 | -0.027 | 0.01 | -0.053 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 17 | 0 | 0.005 | -0.028 | 0.036 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 18 | 0.008 | 0 | -0.028 | 0.01 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 18 | 0 | 0.005 | -0.028 | 0.02 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 19 | 0.008 | 0 | -0.028 | 0.01 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 19 | 0 | 0.005 | -0.028 | 0.009 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 20 | 0.008 | 0 | -0.029 | 0.01 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 20 | 0 | 0.005 | -0.028 | -0.01 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 21 | 0.008 | 0 | -0.03 | 0.01 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 21 | 0 | 0.005 | -0.028 | -0.029 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 22 | 0.008 | 0 | -0.031 | 0.01 | -0.056 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 22 | 0 | 0.005 | -0.028 | -0.064 | -0.056 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 23 | 0.008 | 0 | -0.032 | 0.01 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 23 | 0 | 0.005 | -0.028 | -0.082 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 24 | 0.008 | 0 | -0.033 | 0.01 | -0.055 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 24 | 0 | 0.005 | -0.028 | -0.123 | -0.055 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 25 | 0.008 | 0 | -0.035 | 0.01 | -0.055 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 25 | 0 | 0.005 | -0.028 | -0.16 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 26 | 0.008 | 0 | -0.037 | 0.01 | -0.055 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 26 | 0 | 0.005 | -0.028 | -0.205 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 27 | 0.008 | 0 | -0.039 | 0.01 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 27 | 0 | 0.005 | -0.028 | -0.252 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 28 | 0.008 | 0 | -0.041 | 0.01 | -0.055 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 28 | 0 | 0.005 | -0.028 | -0.309 | -0.055 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 29 | 0.008 | 0 | -0.043 | 0.01 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 29 | 0 | 0.005 | -0.028 | -0.367 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 30 | 0.008 | 0 | -0.046 | 0.01 | -0.055 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 30 | 0 | 0.005 | -0.028 | -0.439 | -0.055 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 31 | 0.008 | 0 | -0.049 | 0.01 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 31 | 0 | 0.005 | -0.028 | -0.506 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 32 | 0.008 | 0 | -0.052 | 0.01 | -0.056 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 32 | 0 | 0.005 | -0.028 | -0.595 | -0.055 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 33 | 0.008 | 0 | -0.055 | 0.01 | -0.055 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 33 | 0 | 0.005 | -0.028 | -0.668 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 34 | 0.008 | 0 | -0.059 | 0.01 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 34 | 0 | 0.005 | -0.028 | -0.753 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 35 | 0.008 | 0 | -0.063 | 0.01 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 35 | 0 | 0.005 | -0.028 | -0.847 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 36 | 0.008 | 0 | -0.067 | 0.01 | -0.053 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 36 | 0 | 0.005 | -0.028 | -0.949 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 37 | 0.008 | 0 | -0.071 | 0.01 | -0.055 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 37 | 0 | 0.005 | -0.028 | -1.057 | -0.055 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 38 | 0.008 | 0 | -0.075 | 0.01 | -0.055 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 38 | 0 | 0.005 | -0.028 | -1.165 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 39 | 0.008 | 0 | -0.08 | 0.01 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 39 | 0 | 0.005 | -0.028 | -1.271 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 40 | 0.008 | 0 | -0.084 | 0.01 | -0.053 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 40 | 0 | 0.005 | -0.028 | -1.383 | -0.054 |  | -0.058 | -0.045 | -0.139
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 15 | 0.008 | 0 | -0.026 | 1.505 | 0.04 |  | -0.058 | -2.657 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 15 | 0 | 0.0001 | -0.028 | 1.941 | -0.041 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 16 | 0.008 | 0 | -0.027 | 1.52 | -0.05 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 16 | 0 | 0.0001 | -0.028 | 1.637 | -0.042 |  | -0.058 | -2.656 | -1.451
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 17 | 0.008 | 0 | -0.027 | 1.512 | 0.011 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 17 | 0 | 0.0001 | -0.028 | 1.654 | -0.039 |  | -0.058 | -2.654 | -1.45
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 18 | 0.008 | 0 | -0.028 | 1.498 | -0.043 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 18 | 0 | 0.0001 | -0.028 | 1.643 | -0.039 |  | -0.058 | -2.653 | -1.449
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 19 | 0.008 | 0 | -0.028 | 1.494 | -0.022 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 19 | 0 | 0.0001 | -0.028 | 1.443 | -0.041 |  | -0.058 | -2.655 | -1.451
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 20 | 0.008 | 0 | -0.029 | 1.511 | -0.047 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 20 | 0 | 0.0001 | -0.028 | 1.32 | -0.042 |  | -0.058 | -2.656 | -1.451
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 21 | 0.008 | 0 | -0.03 | 1.519 | -0.01 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 21 | 0 | 0.0001 | -0.028 | 1.548 | -0.04 |  | -0.058 | -2.654 | -1.449
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 22 | 0.008 | 0 | -0.031 | 1.519 | -0.117 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 22 | 0 | 0.0001 | -0.028 | 0.82 | -0.042 |  | -0.058 | -2.655 | -1.45
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 23 | 0.008 | 0 | -0.032 | 1.516 | -0.012 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 23 | 0 | 0.0001 | -0.028 | 1.46 | -0.041 |  | -0.058 | -2.655 | -1.45
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 24 | 0.008 | 0 | -0.033 | 1.515 | -0.055 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 24 | 0 | 0.0001 | -0.028 | 1.166 | -0.041 |  | -0.058 | -2.655 | -1.451
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 25 | 0.008 | 0 | -0.035 | 1.494 | -0.046 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 25 | 0 | 0.0001 | -0.028 | 1.239 | -0.041 |  | -0.058 | -2.655 | -1.451
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 26 | 0.008 | 0 | -0.037 | 1.523 | -0.048 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 26 | 0 | 0.0001 | -0.028 | 1.106 | -0.041 |  | -0.058 | -2.655 | -1.451
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 27 | 0.008 | 0 | -0.039 | 1.519 | -0.018 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 27 | 0 | 0.0001 | -0.028 | 1.21 | -0.042 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 28 | 0.008 | 0 | -0.041 | 1.516 | -0.067 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 28 | 0 | 0.0001 | -0.028 | 1.01 | -0.042 |  | -0.058 | -2.655 | -1.451
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 29 | 0.008 | 0 | -0.043 | 1.52 | -0.015 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 29 | 0 | 0.0001 | -0.028 | 1.075 | -0.04 |  | -0.058 | -2.654 | -1.45
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 30 | 0.008 | 0 | -0.046 | 1.492 | -0.085 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 30 | 0 | 0.0001 | -0.028 | 0.827 | -0.04 |  | -0.058 | -2.653 | -1.449
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 31 | 0.008 | 0 | -0.049 | 1.51 | -0.018 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 31 | 0 | 0.0001 | -0.028 | 0.978 | -0.04 |  | -0.058 | -2.654 | -1.45
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 32 | 0.008 | 0 | -0.052 | 1.514 | -0.127 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 32 | 0 | 0.0001 | -0.028 | 0.415 | -0.041 |  | -0.058 | -2.654 | -1.45
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 33 | 0.008 | 0 | -0.055 | 1.496 | -0.058 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 33 | 0 | 0.0001 | -0.028 | 0.821 | -0.04 |  | -0.058 | -2.653 | -1.449
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 34 | 0.008 | 0 | -0.059 | 1.496 | -0.007 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 34 | 0 | 0.0001 | -0.028 | 0.532 | -0.041 |  | -0.058 | -2.655 | -1.45
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 35 | 0.008 | 0 | -0.063 | 1.519 | -0.012 |  | -0.058 | -2.657 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 35 | 0 | 0.0001 | -0.028 | 0.692 | -0.04 |  | -0.058 | -2.654 | -1.45
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 36 | 0.008 | 0 | -0.067 | 1.52 | 0.018 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 36 | 0 | 0.0001 | -0.028 | 0.641 | -0.041 |  | -0.058 | -2.655 | -1.451
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 37 | 0.008 | 0 | -0.071 | 1.524 | -0.072 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 37 | 0 | 0.0001 | -0.028 | 0.33 | -0.042 |  | -0.058 | -2.655 | -1.45
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 38 | 0.008 | 0 | -0.075 | 1.499 | -0.098 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 38 | 0 | 0.0001 | -0.028 | 0.257 | -0.042 |  | -0.058 | -2.656 | -1.451
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 39 | 0.008 | 0 | -0.08 | 1.517 | -0.012 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 39 | 0 | 0.0001 | -0.028 | 0.386 | -0.041 |  | -0.058 | -2.656 | -1.451
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 40 | 0.008 | 0 | -0.084 | 1.514 | -0.007 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 40 | 0 | 0.0001 | -0.028 | 0.249 | -0.042 |  | -0.058 | -2.656 | -1.452
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 15 | 0.008 | 0 | -0.026 | 2.151 | 0.157 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 15 | 0 | 5e-05 | -0.028 | 2.266 | -0.024 |  | -0.058 | -12.234 | -2.788 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 16 | 0.008 | 0 | -0.027 | 2.178 | -0.048 |  | -0.058 | -12.234 | -2.788 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 16 | 0 | 5e-05 | -0.028 | 2.248 | -0.025 |  | -0.058 | -12.234 | -2.788 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 17 | 0.008 | 0 | -0.027 | 2.167 | 0.08 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 17 | 0 | 5e-05 | -0.028 | 2.571 | -0.026 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 18 | 0.008 | 0 | -0.028 | 2.156 | 0.001 |  | -0.058 | -12.234 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 18 | 0 | 5e-05 | -0.028 | 1.77 | -0.027 |  | -0.058 | -12.236 | -2.79 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 19 | 0.008 | 0 | -0.028 | 2.129 | 0.009 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 19 | 0 | 5e-05 | -0.028 | 1.968 | -0.028 |  | -0.058 | -12.237 | -2.791 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 20 | 0.008 | 0 | -0.029 | 2.202 | -0.066 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 20 | 0 | 5e-05 | -0.028 | 1.71 | -0.026 |  | -0.058 | -12.232 | -2.787 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 21 | 0.008 | 0 | -0.03 | 2.199 | 0.033 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 21 | 0 | 5e-05 | -0.028 | 1.291 | -0.026 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 22 | 0.008 | 0 | -0.031 | 2.167 | -0.165 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 22 | 0 | 5e-05 | -0.028 | 0.885 | -0.027 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 23 | 0.008 | 0 | -0.032 | 2.16 | 0.045 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 23 | 0 | 5e-05 | -0.028 | 1.507 | -0.025 |  | -0.058 | -12.234 | -2.788 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 24 | 0.008 | 0 | -0.033 | 2.156 | -0.049 |  | -0.058 | -12.234 | -2.788 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 24 | 0 | 5e-05 | -0.028 | 0.5 | -0.024 |  | -0.058 | -12.231 | -2.786 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 25 | 0.008 | 0 | -0.035 | 2.206 | -0.063 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 25 | 0 | 5e-05 | -0.028 | 0.96 | -0.027 |  | -0.058 | -12.236 | -2.79 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 26 | 0.008 | 0 | -0.037 | 2.15 | -0.074 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 26 | 0 | 5e-05 | -0.028 | 0.936 | -0.025 |  | -0.058 | -12.233 | -2.788 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 27 | 0.008 | 0 | -0.039 | 2.181 | -0.003 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 27 | 0 | 5e-05 | -0.028 | 1.861 | -0.027 |  | -0.058 | -12.236 | -2.79 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 28 | 0.008 | 0 | -0.041 | 2.192 | -0.003 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 28 | 0 | 5e-05 | -0.028 | 0.458 | -0.028 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 29 | 0.008 | 0 | -0.043 | 2.168 | 0.032 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 29 | 0 | 5e-05 | -0.028 | 0.752 | -0.028 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 30 | 0.008 | 0 | -0.046 | 2.199 | -0.081 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 30 | 0 | 5e-05 | -0.028 | 0.601 | -0.026 |  | -0.058 | -12.233 | -2.788 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 31 | 0.008 | 0 | -0.049 | 2.192 | 0.046 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 31 | 0 | 5e-05 | -0.028 | 1.567 | -0.027 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 32 | 0.008 | 0 | -0.052 | 2.202 | -0.27 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 32 | 0 | 5e-05 | -0.028 | -0.577 | -0.026 |  | -0.058 | -12.234 | -2.788 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 33 | 0.008 | 0 | -0.055 | 2.154 | -0.048 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 33 | 0 | 5e-05 | -0.028 | 1.173 | -0.027 |  | -0.058 | -12.236 | -2.79 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 34 | 0.008 | 0 | -0.059 | 2.179 | 0.098 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 34 | 0 | 5e-05 | -0.028 | 0.131 | -0.023 |  | -0.058 | -12.232 | -2.787 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 35 | 0.008 | 0 | -0.063 | 2.167 | -0.019 |  | -0.058 | -12.234 | -2.788 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 35 | 0 | 5e-05 | -0.028 | 1.223 | -0.026 |  | -0.058 | -12.236 | -2.79 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 36 | 0.008 | 0 | -0.067 | 2.169 | 0.079 |  | -0.058 | -12.234 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 36 | 0 | 5e-05 | -0.028 | 1.039 | -0.027 |  | -0.058 | -12.236 | -2.79 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 37 | 0.008 | 0 | -0.071 | 2.189 | 0.002 |  | -0.058 | -12.234 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 37 | 0 | 5e-05 | -0.028 | 0.573 | -0.028 |  | -0.058 | -12.236 | -2.791 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 38 | 0.008 | 0 | -0.075 | 2.178 | -0.102 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 38 | 0 | 5e-05 | -0.028 | 0.635 | -0.027 |  | -0.058 | -12.236 | -2.79 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 39 | 0.008 | 0 | -0.08 | 2.158 | -0.052 |  | -0.058 | -12.235 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 39 | 0 | 5e-05 | -0.028 | -0.029 | -0.025 |  | -0.058 | -12.234 | -2.789 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 40 | 0.008 | 0 | -0.084 | 2.166 | 0.057 |  | -0.058 | -12.234 | -2.788 |  | 摸底
case15 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 40 | 0 | 5e-05 | -0.028 | 0.37 | -0.028 |  | -0.058 | -12.237 | -2.791 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 15 | 0.008 | 0 | -0.012 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 15 | 0 | 0.005 | 0.005 | -0.459 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 16 | 0.008 | 0 | -0.019 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 16 | 0 | 0.005 | 0.005 | -0.613 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 17 | 0.008 | 0 | -0.025 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 17 | 0 | 0.005 | 0.005 | -0.782 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 18 | 0.008 | 0 | -0.033 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 18 | 0 | 0.005 | 0.005 | -0.977 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 19 | 0.008 | 0 | -0.042 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 19 | 0 | 0.005 | 0.005 | -1.19 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 20 | 0.008 | 0 | -0.051 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 20 | 0 | 0.005 | 0.005 | -1.414 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 21 | 0.008 | 0 | -0.061 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 21 | 0 | 0.005 | 0.005 | -1.656 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 22 | 0.008 | 0 | -0.071 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 22 | 0 | 0.005 | 0.005 | -1.906 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 23 | 0.008 | 0 | -0.081 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 23 | 0 | 0.005 | 0.005 | -2.165 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 24 | 0.008 | 0 | -0.092 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 24 | 0 | 0.005 | 0.005 | -2.424 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 25 | 0.008 | 0 | -0.102 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 25 | 0 | 0.005 | 0.005 | -2.681 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 26 | 0.008 | 0 | -0.112 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 26 | 0 | 0.005 | 0.005 | -2.933 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 27 | 0.008 | 0 | -0.122 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 27 | 0 | 0.005 | 0.005 | -3.176 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 28 | 0.008 | 0 | -0.131 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 28 | 0 | 0.005 | 0.005 | -3.407 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 29 | 0.008 | 0 | -0.14 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 29 | 0 | 0.005 | 0.005 | -3.623 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 30 | 0.008 | 0 | -0.148 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 30 | 0 | 0.005 | 0.005 | -3.82 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 31 | 0.008 | 0 | -0.155 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 31 | 0 | 0.005 | 0.005 | -4.003 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 32 | 0.008 | 0 | -0.129 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 32 | 0 | 0.005 | 0.005 | -3.347 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 33 | 0.008 | 0 | -0.167 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 33 | 0 | 0.005 | 0.005 | -4.309 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 34 | 0.008 | 0 | -0.172 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 34 | 0 | 0.005 | 0.005 | -4.433 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 35 | 0.008 | 0 | -0.177 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 35 | 0 | 0.005 | 0.005 | -4.542 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 36 | 0.008 | 0 | -0.18 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 36 | 0 | 0.005 | 0.005 | -4.633 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 37 | 0.008 | 0 | -0.184 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 37 | 0 | 0.005 | 0.005 | -4.709 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 38 | 0.008 | 0 | -0.186 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 38 | 0 | 0.005 | 0.005 | -4.771 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 39 | 0.008 | 0 | -0.188 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 39 | 0 | 0.005 | 0.005 | -4.821 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 40 | 0.008 | 0 | -0.19 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 40 | 0 | 0.005 | 0.005 | -4.862 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 15 | 0.008 | 0 | -0.012 | -1.161 | 0.074 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 15 | 0 | 0.0001 | 0.005 | -1.423 | 0.043 |  | 0.253 | -1.241 | 0.785
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 16 | 0.008 | 0 | -0.019 | -1.16 | 0.033 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 16 | 0 | 0.0001 | 0.005 | -1.782 | 0.041 |  | 0.253 | -1.242 | 0.783
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 17 | 0.008 | 0 | -0.025 | -1.162 | 0.062 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 17 | 0 | 0.0001 | 0.005 | -1.845 | 0.042 |  | 0.253 | -1.241 | 0.785
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 18 | 0.008 | 0 | -0.033 | -1.161 | 0.039 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 18 | 0 | 0.0001 | 0.005 | -2.123 | 0.042 |  | 0.253 | -1.241 | 0.785
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 19 | 0.008 | 0 | -0.042 | -1.161 | 0.03 |  | 0.253 | -1.241 | 0.785
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 19 | 0 | 0.0001 | 0.005 | -2.378 | 0.042 |  | 0.253 | -1.241 | 0.785
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 20 | 0.008 | 0 | -0.051 | -1.162 | 0.07 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 20 | 0 | 0.0001 | 0.005 | -2.423 | 0.041 |  | 0.253 | -1.242 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 21 | 0.008 | 0 | -0.061 | -1.16 | 0.056 |  | 0.253 | -1.242 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 21 | 0 | 0.0001 | 0.005 | -2.728 | 0.043 |  | 0.253 | -1.241 | 0.785
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 22 | 0.008 | 0 | -0.071 | -1.161 | 0.057 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 22 | 0 | 0.0001 | 0.005 | -2.943 | 0.042 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 23 | 0.008 | 0 | -0.081 | -1.16 | 0.038 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 23 | 0 | 0.0001 | 0.005 | -3.295 | 0.042 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 24 | 0.008 | 0 | -0.092 | -1.16 | 0.033 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 24 | 0 | 0.0001 | 0.005 | -3.61 | 0.041 |  | 0.253 | -1.242 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 25 | 0.008 | 0 | -0.102 | -1.159 | 0.055 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 25 | 0 | 0.0001 | 0.005 | -3.762 | 0.042 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 26 | 0.008 | 0 | -0.112 | -1.162 | 0.051 |  | 0.253 | -1.241 | 0.785
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 26 | 0 | 0.0001 | 0.005 | -4.018 | 0.042 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 27 | 0.008 | 0 | -0.122 | -1.162 | 0.036 |  | 0.253 | -1.241 | 0.785
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 27 | 0 | 0.0001 | 0.005 | -4.369 | 0.042 |  | 0.253 | -1.242 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 28 | 0.008 | 0 | -0.131 | -1.16 | 0.039 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 28 | 0 | 0.0001 | 0.005 | -4.575 | 0.041 |  | 0.253 | -1.242 | 0.783
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 29 | 0.008 | 0 | -0.14 | -1.161 | 0.02 |  | 0.253 | -1.241 | 0.785
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 29 | 0 | 0.0001 | 0.005 | -4.85 | 0.042 |  | 0.253 | -1.242 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 30 | 0.008 | 0 | -0.148 | -1.159 | 0.043 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 30 | 0 | 0.0001 | 0.005 | -4.938 | 0.042 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 31 | 0.008 | 0 | -0.155 | -1.161 | 0.031 |  | 0.253 | -1.242 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 31 | 0 | 0.0001 | 0.005 | -5.158 | 0.041 |  | 0.253 | -1.242 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 32 | 0.008 | 0 | -0.129 | -1.16 | 0.05 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 32 | 0 | 0.0001 | 0.005 | -4.422 | 0.041 |  | 0.253 | -1.242 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 33 | 0.008 | 0 | -0.167 | -1.158 | 0.037 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 33 | 0 | 0.0001 | 0.005 | -5.486 | 0.041 |  | 0.253 | -1.242 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 34 | 0.008 | 0 | -0.172 | -1.162 | 0.043 |  | 0.253 | -1.241 | 0.785
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 34 | 0 | 0.0001 | 0.005 | -5.564 | 0.041 |  | 0.253 | -1.242 | 0.783
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 35 | 0.008 | 0 | -0.177 | -1.158 | 0.03 |  | 0.253 | -1.241 | 0.785
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 35 | 0 | 0.0001 | 0.005 | -5.731 | 0.042 |  | 0.253 | -1.241 | 0.785
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 36 | 0.008 | 0 | -0.18 | -1.154 | 0.039 |  | 0.253 | -1.242 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 36 | 0 | 0.0001 | 0.005 | -5.791 | 0.041 |  | 0.253 | -1.242 | 0.783
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 37 | 0.008 | 0 | -0.184 | -1.159 | 0.04 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 37 | 0 | 0.0001 | 0.005 | -5.88 | 0.043 |  | 0.253 | -1.24 | 0.785
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 38 | 0.008 | 0 | -0.186 | -1.161 | 0.046 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 38 | 0 | 0.0001 | 0.005 | -5.878 | 0.041 |  | 0.253 | -1.242 | 0.783
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 39 | 0.008 | 0 | -0.188 | -1.16 | 0.045 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 39 | 0 | 0.0001 | 0.005 | -5.944 | 0.042 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 40 | 0.008 | 0 | -0.19 | -1.161 | 0.04 |  | 0.253 | -1.241 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 40 | 0 | 0.0001 | 0.005 | -6.005 | 0.041 |  | 0.253 | -1.242 | 0.784
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 15 | 0.008 | 0 | -0.012 | -1.995 | 0.122 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 15 | 0 | 5e-05 | 0.005 | -2.129 | 0.056 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 16 | 0.008 | 0 | -0.019 | -1.994 | 0.039 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 16 | 0 | 5e-05 | 0.005 | -2.543 | 0.054 |  | 0.253 | -3.524 | 1.046 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 17 | 0.008 | 0 | -0.025 | -1.99 | 0.077 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 17 | 0 | 5e-05 | 0.005 | -2.614 | 0.052 |  | 0.253 | -3.526 | 1.043 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 18 | 0.008 | 0 | -0.033 | -1.983 | 0.07 |  | 0.253 | -3.523 | 1.046 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 18 | 0 | 5e-05 | 0.005 | -2.897 | 0.053 |  | 0.253 | -3.525 | 1.045 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 19 | 0.008 | 0 | -0.042 | -1.99 | 0.032 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 19 | 0 | 5e-05 | 0.005 | -3.241 | 0.055 |  | 0.253 | -3.523 | 1.046 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 20 | 0.008 | 0 | -0.051 | -1.99 | 0.1 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 20 | 0 | 5e-05 | 0.005 | -3.116 | 0.056 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 21 | 0.008 | 0 | -0.061 | -1.995 | 0.086 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 21 | 0 | 5e-05 | 0.005 | -3.457 | 0.055 |  | 0.253 | -3.523 | 1.046 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 22 | 0.008 | 0 | -0.071 | -1.992 | 0.087 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 22 | 0 | 5e-05 | 0.005 | -3.672 | 0.054 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 23 | 0.008 | 0 | -0.081 | -1.998 | 0.065 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 23 | 0 | 5e-05 | 0.005 | -4.013 | 0.055 |  | 0.253 | -3.524 | 1.046 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 24 | 0.008 | 0 | -0.092 | -1.992 | 0.052 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 24 | 0 | 5e-05 | 0.005 | -4.448 | 0.054 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 25 | 0.008 | 0 | -0.102 | -1.994 | 0.07 |  | 0.253 | -3.523 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 25 | 0 | 5e-05 | 0.005 | -4.57 | 0.054 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 26 | 0.008 | 0 | -0.112 | -1.994 | 0.069 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 26 | 0 | 5e-05 | 0.005 | -4.849 | 0.057 |  | 0.253 | -3.522 | 1.048 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 27 | 0.008 | 0 | -0.122 | -1.997 | 0.048 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 27 | 0 | 5e-05 | 0.005 | -5.215 | 0.054 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 28 | 0.008 | 0 | -0.131 | -1.996 | 0.041 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 28 | 0 | 5e-05 | 0.005 | -5.379 | 0.057 |  | 0.253 | -3.521 | 1.048 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 29 | 0.008 | 0 | -0.14 | -1.998 | 0.022 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 29 | 0 | 5e-05 | 0.005 | -5.72 | 0.056 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 30 | 0.008 | 0 | -0.148 | -1.996 | 0.048 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 30 | 0 | 5e-05 | 0.005 | -5.842 | 0.055 |  | 0.253 | -3.523 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 31 | 0.008 | 0 | -0.155 | -1.99 | 0.04 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 31 | 0 | 5e-05 | 0.005 | -6.115 | 0.054 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 32 | 0.008 | 0 | -0.129 | -1.99 | 0.08 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 32 | 0 | 5e-05 | 0.005 | -5.148 | 0.054 |  | 0.253 | -3.525 | 1.045 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 33 | 0.008 | 0 | -0.167 | -1.981 | 0.04 |  | 0.253 | -3.523 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 33 | 0 | 5e-05 | 0.005 | -6.325 | 0.055 |  | 0.253 | -3.523 | 1.046 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 34 | 0.008 | 0 | -0.172 | -1.996 | 0.05 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 34 | 0 | 5e-05 | 0.005 | -6.364 | 0.055 |  | 0.253 | -3.523 | 1.046 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 35 | 0.008 | 0 | -0.177 | -1.991 | 0.036 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 35 | 0 | 5e-05 | 0.005 | -6.569 | 0.054 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 36 | 0.008 | 0 | -0.18 | -1.995 | 0.048 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 36 | 0 | 5e-05 | 0.005 | -6.61 | 0.054 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 37 | 0.008 | 0 | -0.184 | -1.994 | 0.052 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 37 | 0 | 5e-05 | 0.005 | -6.704 | 0.055 |  | 0.253 | -3.523 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 38 | 0.008 | 0 | -0.186 | -1.996 | 0.06 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 38 | 0 | 5e-05 | 0.005 | -6.683 | 0.053 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 39 | 0.008 | 0 | -0.188 | -1.996 | 0.06 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 39 | 0 | 5e-05 | 0.005 | -6.829 | 0.053 |  | 0.253 | -3.525 | 1.045 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 40 | 0.008 | 0 | -0.19 | -1.992 | 0.055 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 40 | 0 | 5e-05 | 0.005 | -6.799 | 0.056 |  | 0.253 | -3.523 | 1.047 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 15 | 0.008 | 0 | -0.026 | 0.011 | -0.052 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 15 | 0 | 0.005 | -0.028 | 0.051 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 16 | 0.008 | 0 | -0.027 | 0.011 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 16 | 0 | 0.005 | -0.028 | 0.037 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 17 | 0.008 | 0 | -0.027 | 0.011 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 17 | 0 | 0.005 | -0.028 | 0.036 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 18 | 0.008 | 0 | -0.027 | 0.011 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 18 | 0 | 0.005 | -0.028 | 0.02 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 19 | 0.008 | 0 | -0.028 | 0.011 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 19 | 0 | 0.005 | -0.028 | 0.009 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 20 | 0.008 | 0 | -0.028 | 0.011 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 20 | 0 | 0.005 | -0.028 | -0.011 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 21 | 0.008 | 0 | -0.029 | 0.011 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 21 | 0 | 0.005 | -0.028 | -0.029 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 22 | 0.008 | 0 | -0.03 | 0.011 | -0.055 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 22 | 0 | 0.005 | -0.028 | -0.065 | -0.055 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 23 | 0.008 | 0 | -0.031 | 0.011 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 23 | 0 | 0.005 | -0.028 | -0.082 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 24 | 0.008 | 0 | -0.033 | 0.011 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 24 | 0 | 0.005 | -0.028 | -0.122 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 25 | 0.008 | 0 | -0.035 | 0.011 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 25 | 0 | 0.005 | -0.028 | -0.159 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 26 | 0.008 | 0 | -0.036 | 0.011 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 26 | 0 | 0.005 | -0.028 | -0.204 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 27 | 0.008 | 0 | -0.038 | 0.011 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 27 | 0 | 0.005 | -0.028 | -0.249 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 28 | 0.008 | 0 | -0.041 | 0.011 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 28 | 0 | 0.005 | -0.028 | -0.309 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 29 | 0.008 | 0 | -0.043 | 0.011 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 29 | 0 | 0.005 | -0.028 | -0.365 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 30 | 0.008 | 0 | -0.046 | 0.011 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 30 | 0 | 0.005 | -0.028 | -0.437 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 31 | 0.008 | 0 | -0.048 | 0.011 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 31 | 0 | 0.005 | -0.028 | -0.502 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 32 | 0.008 | 0 | -0.052 | 0.011 | -0.056 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 32 | 0 | 0.005 | -0.028 | -0.594 | -0.055 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 33 | 0.008 | 0 | -0.055 | 0.011 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 33 | 0 | 0.005 | -0.028 | -0.668 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 34 | 0.008 | 0 | -0.058 | 0.011 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 34 | 0 | 0.005 | -0.028 | -0.755 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 35 | 0.008 | 0 | -0.062 | 0.011 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 35 | 0 | 0.005 | -0.028 | -0.849 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 36 | 0.008 | 0 | -0.066 | 0.011 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 36 | 0 | 0.005 | -0.028 | -0.947 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 37 | 0.008 | 0 | -0.071 | 0.011 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 37 | 0 | 0.005 | -0.028 | -1.053 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 38 | 0.008 | 0 | -0.075 | 0.011 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 38 | 0 | 0.005 | -0.028 | -1.164 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 39 | 0.008 | 0 | -0.079 | 0.011 | -0.053 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 39 | 0 | 0.005 | -0.028 | -1.273 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 40 | 0.008 | 0 | -0.084 | 0.011 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 40 | 0 | 0.005 | -0.028 | -1.386 | -0.054 |  | -0.059 | -0.046 | -0.141
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 15 | 0.008 | 0 | -0.026 | 1.461 | 0.049 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 15 | 0 | 0.0001 | -0.028 | 1.962 | -0.041 |  | -0.059 | -2.658 | -1.454
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 16 | 0.008 | 0 | -0.027 | 1.469 | -0.055 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 16 | 0 | 0.0001 | -0.028 | 1.504 | -0.041 |  | -0.059 | -2.656 | -1.453
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 17 | 0.008 | 0 | -0.027 | 1.452 | 0.011 |  | -0.059 | -2.655 | -1.451
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 17 | 0 | 0.0001 | -0.028 | 1.775 | -0.04 |  | -0.059 | -2.656 | -1.453
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 18 | 0.008 | 0 | -0.027 | 1.461 | -0.031 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 18 | 0 | 0.0001 | -0.028 | 1.575 | -0.039 |  | -0.059 | -2.655 | -1.451
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 19 | 0.008 | 0 | -0.028 | 1.468 | -0.015 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 19 | 0 | 0.0001 | -0.028 | 1.516 | -0.04 |  | -0.059 | -2.656 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 20 | 0.008 | 0 | -0.028 | 1.469 | -0.037 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 20 | 0 | 0.0001 | -0.028 | 1.422 | -0.041 |  | -0.059 | -2.656 | -1.453
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 21 | 0.008 | 0 | -0.029 | 1.465 | -0.014 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 21 | 0 | 0.0001 | -0.028 | 1.592 | -0.041 |  | -0.059 | -2.656 | -1.453
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 22 | 0.008 | 0 | -0.031 | 1.467 | -0.119 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 22 | 0 | 0.0001 | -0.028 | 0.888 | -0.041 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 23 | 0.008 | 0 | -0.031 | 1.463 | -0.002 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 23 | 0 | 0.0001 | -0.028 | 1.726 | -0.039 |  | -0.059 | -2.656 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 24 | 0.008 | 0 | -0.033 | 1.467 | -0.065 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 24 | 0 | 0.0001 | -0.028 | 1.266 | -0.04 |  | -0.059 | -2.656 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 25 | 0.008 | 0 | -0.035 | 1.467 | -0.038 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 25 | 0 | 0.0001 | -0.028 | 1.194 | -0.041 |  | -0.059 | -2.657 | -1.453
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 26 | 0.008 | 0 | -0.036 | 1.471 | -0.042 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 26 | 0 | 0.0001 | -0.028 | 1.167 | -0.039 |  | -0.059 | -2.654 | -1.451
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 27 | 0.008 | 0 | -0.038 | 1.473 | -0.025 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 27 | 0 | 0.0001 | -0.028 | 1.344 | -0.04 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 28 | 0.008 | 0 | -0.041 | 1.465 | -0.049 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 28 | 0 | 0.0001 | -0.028 | 0.983 | -0.039 |  | -0.059 | -2.655 | -1.451
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 29 | 0.008 | 0 | -0.043 | 1.468 | -0.008 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 29 | 0 | 0.0001 | -0.028 | 1.262 | -0.04 |  | -0.059 | -2.656 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 30 | 0.008 | 0 | -0.046 | 1.466 | -0.088 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 30 | 0 | 0.0001 | -0.028 | 0.703 | -0.04 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 31 | 0.008 | 0 | -0.048 | 1.46 | -0.02 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 31 | 0 | 0.0001 | -0.028 | 1.06 | -0.04 |  | -0.059 | -2.656 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 32 | 0.008 | 0 | -0.052 | 1.469 | -0.152 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 32 | 0 | 0.0001 | -0.028 | 0.492 | -0.041 |  | -0.059 | -2.656 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 33 | 0.008 | 0 | -0.055 | 1.477 | -0.029 |  | -0.059 | -2.655 | -1.451
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 33 | 0 | 0.0001 | -0.028 | 0.835 | -0.041 |  | -0.059 | -2.656 | -1.453
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 34 | 0.008 | 0 | -0.059 | 1.467 | -0.014 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 34 | 0 | 0.0001 | -0.028 | 0.752 | -0.041 |  | -0.059 | -2.657 | -1.454
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 35 | 0.008 | 0 | -0.062 | 1.46 | -0.009 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 35 | 0 | 0.0001 | -0.028 | 0.699 | -0.039 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 36 | 0.008 | 0 | -0.066 | 1.465 | 0 |  | -0.059 | -2.656 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 36 | 0 | 0.0001 | -0.028 | 0.486 | -0.041 |  | -0.059 | -2.656 | -1.453
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 37 | 0.008 | 0 | -0.071 | 1.464 | -0.097 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 37 | 0 | 0.0001 | -0.028 | 0.02 | -0.04 |  | -0.059 | -2.656 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 38 | 0.008 | 0 | -0.075 | 1.459 | -0.063 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 38 | 0 | 0.0001 | -0.028 | 0.133 | -0.04 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 39 | 0.008 | 0 | -0.079 | 1.475 | -0.054 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 39 | 0 | 0.0001 | -0.028 | 0.122 | -0.04 |  | -0.059 | -2.656 | -1.453
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 40 | 0.008 | 0 | -0.084 | 1.463 | -0.034 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 40 | 0 | 0.0001 | -0.028 | 0.096 | -0.039 |  | -0.059 | -2.655 | -1.452
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 15 | 0.008 | 0 | -0.026 | 1.929 | 0.147 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 15 | 0 | 5e-05 | -0.028 | 1.842 | -0.023 |  | -0.059 | -12.233 | -2.789 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 16 | 0.008 | 0 | -0.027 | 1.976 | -0.046 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 16 | 0 | 5e-05 | -0.028 | 1.483 | -0.023 |  | -0.059 | -12.234 | -2.789 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 17 | 0.008 | 0 | -0.027 | 2.052 | 0.068 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 17 | 0 | 5e-05 | -0.028 | 1.542 | -0.024 |  | -0.059 | -12.234 | -2.789 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 18 | 0.008 | 0 | -0.027 | 1.978 | -0.009 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 18 | 0 | 5e-05 | -0.028 | 1.884 | -0.029 |  | -0.059 | -12.239 | -2.794 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 19 | 0.008 | 0 | -0.028 | 1.964 | 0.02 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 19 | 0 | 5e-05 | -0.028 | 2.287 | -0.025 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 20 | 0.008 | 0 | -0.028 | 2.054 | -0.071 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 20 | 0 | 5e-05 | -0.028 | 0.766 | -0.023 |  | -0.059 | -12.234 | -2.789 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 21 | 0.008 | 0 | -0.029 | 2.072 | 0.043 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 21 | 0 | 5e-05 | -0.028 | 1.493 | -0.026 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 22 | 0.008 | 0 | -0.031 | 2.004 | -0.16 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 22 | 0 | 5e-05 | -0.028 | 0.46 | -0.026 |  | -0.059 | -12.235 | -2.79 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 23 | 0.008 | 0 | -0.031 | 2.095 | 0.045 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 23 | 0 | 5e-05 | -0.028 | 1.555 | -0.025 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 24 | 0.008 | 0 | -0.033 | 2.065 | -0.052 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 24 | 0 | 5e-05 | -0.028 | 1.017 | -0.026 |  | -0.059 | -12.234 | -2.79 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 25 | 0.008 | 0 | -0.034 | 2.111 | -0.039 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 25 | 0 | 5e-05 | -0.028 | 1.172 | -0.023 |  | -0.059 | -12.232 | -2.787 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 26 | 0.008 | 0 | -0.036 | 2.075 | -0.058 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 26 | 0 | 5e-05 | -0.028 | 1.134 | -0.026 |  | -0.059 | -12.237 | -2.792 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 27 | 0.008 | 0 | -0.038 | 1.972 | 0.045 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 27 | 0 | 5e-05 | -0.028 | 2.411 | -0.029 |  | -0.059 | -12.239 | -2.794 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 28 | 0.008 | 0 | -0.041 | 1.927 | -0.052 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 28 | 0 | 5e-05 | -0.028 | 1.041 | -0.023 |  | -0.059 | -12.233 | -2.788 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 29 | 0.008 | 0 | -0.043 | 2.038 | 0.03 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 29 | 0 | 5e-05 | -0.028 | 1.504 | -0.026 |  | -0.059 | -12.237 | -2.792 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 30 | 0.008 | 0 | -0.046 | 2.041 | -0.08 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 30 | 0 | 5e-05 | -0.028 | 1.034 | -0.026 |  | -0.059 | -12.237 | -2.792 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 31 | 0.008 | 0 | -0.049 | 2.159 | 0.037 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 31 | 0 | 5e-05 | -0.028 | 0.597 | -0.025 |  | -0.059 | -12.235 | -2.79 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 32 | 0.008 | 0 | -0.052 | 2.077 | -0.235 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 32 | 0 | 5e-05 | -0.028 | -0.662 | -0.028 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 33 | 0.008 | 0 | -0.055 | 2.09 | -0.04 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 33 | 0 | 5e-05 | -0.028 | 1.012 | -0.026 |  | -0.059 | -12.237 | -2.792 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 34 | 0.008 | 0 | -0.059 | 2.005 | 0.031 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 34 | 0 | 5e-05 | -0.028 | 0.479 | -0.025 |  | -0.059 | -12.234 | -2.789 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 35 | 0.008 | 0 | -0.062 | 2.069 | -0.019 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 35 | 0 | 5e-05 | -0.028 | 0.079 | -0.025 |  | -0.059 | -12.233 | -2.788 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 36 | 0.008 | 0 | -0.066 | 1.926 | -0.003 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 36 | 0 | 5e-05 | -0.028 | 0.823 | -0.025 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 37 | 0.008 | 0 | -0.071 | 2.065 | 0.011 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 37 | 0 | 5e-05 | -0.028 | 0.656 | -0.023 |  | -0.059 | -12.234 | -2.789 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 38 | 0.008 | 0 | -0.076 | 2.02 | -0.084 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 38 | 0 | 5e-05 | -0.028 | 0.625 | -0.028 |  | -0.059 | -12.239 | -2.794 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 39 | 0.008 | 0 | -0.079 | 2.061 | -0.128 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 39 | 0 | 5e-05 | -0.028 | -0.603 | -0.025 |  | -0.059 | -12.233 | -2.788 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 40 | 0.008 | 0 | -0.084 | 2.055 | 0.114 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 40 | 0 | 5e-05 | -0.028 | 0.767 | -0.026 |  | -0.059 | -12.236 | -2.791 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 15 | 0.008 | 0 | -0.012 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 15 | 0 | 0.005 | 0.005 | -0.459 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 16 | 0.008 | 0 | -0.019 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 16 | 0 | 0.005 | 0.005 | -0.613 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 17 | 0.008 | 0 | -0.025 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 17 | 0 | 0.005 | 0.005 | -0.783 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 18 | 0.008 | 0 | -0.033 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 18 | 0 | 0.005 | 0.005 | -0.976 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 19 | 0.008 | 0 | -0.042 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 19 | 0 | 0.005 | 0.005 | -1.189 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 20 | 0.008 | 0 | -0.051 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 20 | 0 | 0.005 | 0.005 | -1.413 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 21 | 0.008 | 0 | -0.061 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 21 | 0 | 0.005 | 0.005 | -1.656 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 22 | 0.008 | 0 | -0.071 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 22 | 0 | 0.005 | 0.005 | -1.906 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 23 | 0.008 | 0 | -0.081 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 23 | 0 | 0.005 | 0.005 | -2.165 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 24 | 0.008 | 0 | -0.092 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 24 | 0 | 0.005 | 0.005 | -2.424 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 25 | 0.008 | 0 | -0.102 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 25 | 0 | 0.005 | 0.005 | -2.68 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 26 | 0.008 | 0 | -0.112 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 26 | 0 | 0.005 | 0.005 | -2.933 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 27 | 0.008 | 0 | -0.122 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 27 | 0 | 0.005 | 0.005 | -3.177 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 28 | 0.008 | 0 | -0.131 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 28 | 0 | 0.005 | 0.005 | -3.407 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 29 | 0.008 | 0 | -0.14 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 29 | 0 | 0.005 | 0.005 | -3.623 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 30 | 0.008 | 0 | -0.148 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 30 | 0 | 0.005 | 0.005 | -3.82 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 31 | 0.008 | 0 | -0.155 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 31 | 0 | 0.005 | 0.005 | -4.002 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 32 | 0.008 | 0 | -0.129 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 32 | 0 | 0.005 | 0.005 | -3.347 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 33 | 0.008 | 0 | -0.167 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 33 | 0 | 0.005 | 0.005 | -4.309 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 34 | 0.008 | 0 | -0.172 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 34 | 0 | 0.005 | 0.005 | -4.433 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 35 | 0.008 | 0 | -0.177 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 35 | 0 | 0.005 | 0.005 | -4.542 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 36 | 0.008 | 0 | -0.18 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 36 | 0 | 0.005 | 0.005 | -4.633 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 37 | 0.008 | 0 | -0.183 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 37 | 0 | 0.005 | 0.005 | -4.709 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 38 | 0.008 | 0 | -0.186 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 38 | 0 | 0.005 | 0.005 | -4.77 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 39 | 0.008 | 0 | -0.188 | -0.018 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 39 | 0 | 0.005 | 0.005 | -4.821 | 0.029 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 40 | 0.008 | 0 | -0.19 | -0.018 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 40 | 0 | 0.005 | 0.005 | -4.862 | 0.028 |  | 0.253 | 0.233 | 0.527
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 15 | 0.008 | 0 | -0.012 | -1.163 | 0.073 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 15 | 0 | 0.0001 | 0.005 | -1.444 | 0.042 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 16 | 0.008 | 0 | -0.018 | -1.166 | 0.037 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 16 | 0 | 0.0001 | 0.005 | -1.775 | 0.042 |  | 0.253 | -1.241 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 17 | 0.008 | 0 | -0.025 | -1.163 | 0.054 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 17 | 0 | 0.0001 | 0.005 | -1.863 | 0.042 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 18 | 0.008 | 0 | -0.033 | -1.162 | 0.038 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 18 | 0 | 0.0001 | 0.005 | -2.127 | 0.041 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 19 | 0.008 | 0 | -0.042 | -1.158 | 0.03 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 19 | 0 | 0.0001 | 0.005 | -2.372 | 0.042 |  | 0.253 | -1.241 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 20 | 0.008 | 0 | -0.051 | -1.163 | 0.073 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 20 | 0 | 0.0001 | 0.005 | -2.431 | 0.043 |  | 0.253 | -1.241 | 0.785
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 21 | 0.008 | 0 | -0.061 | -1.167 | 0.048 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 21 | 0 | 0.0001 | 0.005 | -2.688 | 0.042 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 22 | 0.008 | 0 | -0.071 | -1.159 | 0.059 |  | 0.253 | -1.243 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 22 | 0 | 0.0001 | 0.005 | -2.959 | 0.041 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 23 | 0.008 | 0 | -0.081 | -1.162 | 0.043 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 23 | 0 | 0.0001 | 0.005 | -3.27 | 0.041 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 24 | 0.008 | 0 | -0.092 | -1.162 | 0.033 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 24 | 0 | 0.0001 | 0.005 | -3.621 | 0.041 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 25 | 0.008 | 0 | -0.102 | -1.163 | 0.055 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 25 | 0 | 0.0001 | 0.005 | -3.741 | 0.042 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 26 | 0.008 | 0 | -0.112 | -1.167 | 0.049 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 26 | 0 | 0.0001 | 0.005 | -4.046 | 0.042 |  | 0.253 | -1.241 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 27 | 0.008 | 0 | -0.122 | -1.164 | 0.031 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 27 | 0 | 0.0001 | 0.005 | -4.348 | 0.04 |  | 0.253 | -1.243 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 28 | 0.008 | 0 | -0.131 | -1.166 | 0.038 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 28 | 0 | 0.0001 | 0.005 | -4.567 | 0.042 |  | 0.253 | -1.241 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 29 | 0.008 | 0 | -0.14 | -1.163 | 0.024 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 29 | 0 | 0.0001 | 0.005 | -4.841 | 0.041 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 30 | 0.008 | 0 | -0.148 | -1.165 | 0.04 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 30 | 0 | 0.0001 | 0.005 | -4.95 | 0.042 |  | 0.253 | -1.241 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 31 | 0.008 | 0 | -0.155 | -1.163 | 0.034 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 31 | 0 | 0.0001 | 0.005 | -5.168 | 0.041 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 32 | 0.008 | 0 | -0.129 | -1.163 | 0.05 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 32 | 0 | 0.0001 | 0.005 | -4.454 | 0.042 |  | 0.253 | -1.241 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 33 | 0.008 | 0 | -0.167 | -1.164 | 0.035 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 33 | 0 | 0.0001 | 0.005 | -5.487 | 0.042 |  | 0.253 | -1.241 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 34 | 0.008 | 0 | -0.172 | -1.163 | 0.04 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 34 | 0 | 0.0001 | 0.005 | -5.573 | 0.042 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 35 | 0.008 | 0 | -0.177 | -1.163 | 0.031 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 35 | 0 | 0.0001 | 0.005 | -5.721 | 0.042 |  | 0.253 | -1.241 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 36 | 0.008 | 0 | -0.18 | -1.162 | 0.038 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 36 | 0 | 0.0001 | 0.005 | -5.775 | 0.042 |  | 0.253 | -1.241 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 37 | 0.008 | 0 | -0.183 | -1.163 | 0.038 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 37 | 0 | 0.0001 | 0.005 | -5.858 | 0.042 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 38 | 0.008 | 0 | -0.186 | -1.162 | 0.042 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 38 | 0 | 0.0001 | 0.005 | -5.888 | 0.042 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 39 | 0.008 | 0 | -0.188 | -1.161 | 0.044 |  | 0.253 | -1.243 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 39 | 0 | 0.0001 | 0.005 | -5.923 | 0.042 |  | 0.253 | -1.242 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 40 | 0.008 | 0 | -0.19 | -1.163 | 0.04 |  | 0.253 | -1.242 | 0.783
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 40 | 0 | 0.0001 | 0.005 | -6.017 | 0.042 |  | 0.253 | -1.241 | 0.784
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 15 | 0.008 | 0 | -0.012 | -1.99 | 0.119 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 15 | 0 | 5e-05 | 0.005 | -2.062 | 0.055 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 16 | 0.008 | 0 | -0.018 | -1.989 | 0.041 |  | 0.253 | -3.525 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 16 | 0 | 5e-05 | 0.005 | -2.643 | 0.054 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 17 | 0.008 | 0 | -0.025 | -1.989 | 0.077 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 17 | 0 | 5e-05 | 0.005 | -2.653 | 0.055 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 18 | 0.008 | 0 | -0.033 | -1.991 | 0.051 |  | 0.253 | -3.525 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 18 | 0 | 5e-05 | 0.005 | -2.944 | 0.056 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 19 | 0.008 | 0 | -0.042 | -1.991 | 0.025 |  | 0.253 | -3.525 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 19 | 0 | 5e-05 | 0.005 | -3.225 | 0.056 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 20 | 0.008 | 0 | -0.051 | -1.988 | 0.104 |  | 0.253 | -3.525 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 20 | 0 | 5e-05 | 0.005 | -3.103 | 0.056 |  | 0.253 | -3.523 | 1.047 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 21 | 0.008 | 0 | -0.061 | -1.986 | 0.086 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 21 | 0 | 5e-05 | 0.005 | -3.486 | 0.055 |  | 0.253 | -3.524 | 1.046 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 22 | 0.008 | 0 | -0.071 | -1.988 | 0.093 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 22 | 0 | 5e-05 | 0.005 | -3.697 | 0.055 |  | 0.253 | -3.523 | 1.046 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 23 | 0.008 | 0 | -0.081 | -1.992 | 0.067 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 23 | 0 | 5e-05 | 0.005 | -4.108 | 0.055 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 24 | 0.008 | 0 | -0.092 | -1.986 | 0.034 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 24 | 0 | 5e-05 | 0.005 | -4.389 | 0.054 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 25 | 0.008 | 0 | -0.102 | -1.989 | 0.062 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 25 | 0 | 5e-05 | 0.005 | -4.664 | 0.054 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 26 | 0.008 | 0 | -0.112 | -1.986 | 0.07 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 26 | 0 | 5e-05 | 0.005 | -4.776 | 0.053 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 27 | 0.008 | 0 | -0.122 | -1.986 | 0.032 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 27 | 0 | 5e-05 | 0.005 | -5.234 | 0.053 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 28 | 0.008 | 0 | -0.131 | -1.986 | 0.051 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 28 | 0 | 5e-05 | 0.005 | -5.419 | 0.054 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 29 | 0.008 | 0 | -0.14 | -1.987 | 0.031 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 29 | 0 | 5e-05 | 0.005 | -5.721 | 0.053 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 30 | 0.008 | 0 | -0.148 | -1.986 | 0.046 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 30 | 0 | 5e-05 | 0.005 | -5.831 | 0.052 |  | 0.253 | -3.526 | 1.043 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 31 | 0.008 | 0 | -0.155 | -1.991 | 0.032 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 31 | 0 | 5e-05 | 0.005 | -6.085 | 0.054 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 32 | 0.008 | 0 | -0.129 | -1.989 | 0.087 |  | 0.253 | -3.525 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 32 | 0 | 5e-05 | 0.005 | -5.094 | 0.054 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 33 | 0.008 | 0 | -0.167 | -1.165 | 0.747 |  | 0.253 | -2.839 | 1.721 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 33 | 0 | 5e-05 | 0.005 | -6.315 | 0.054 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 34 | 0.008 | 0 | -0.172 | -1.987 | 0.049 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 34 | 0 | 5e-05 | 0.005 | -6.37 | 0.055 |  | 0.253 | -3.523 | 1.046 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 35 | 0.008 | 0 | -0.177 | -1.99 | 0.041 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 35 | 0 | 5e-05 | 0.005 | -6.585 | 0.055 |  | 0.253 | -3.523 | 1.046 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 36 | 0.008 | 0 | -0.18 | -1.99 | 0.052 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 36 | 0 | 5e-05 | 0.005 | -6.652 | 0.056 |  | 0.253 | -3.522 | 1.047 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 37 | 0.008 | 0 | -0.183 | -1.992 | 0.046 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 37 | 0 | 5e-05 | 0.005 | -6.716 | 0.055 |  | 0.253 | -3.523 | 1.046 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 38 | 0.008 | 0 | -0.186 | -1.986 | 0.059 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 38 | 0 | 5e-05 | 0.005 | -6.668 | 0.053 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 39 | 0.008 | 0 | -0.188 | -1.988 | 0.056 |  | 0.253 | -3.525 | 1.044 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 39 | 0 | 5e-05 | 0.005 | -6.799 | 0.055 |  | 0.253 | -3.524 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 40 | 0.008 | 0 | -0.19 | -1.99 | 0.05 |  | 0.253 | -3.525 | 1.045 |  | 摸底
case15 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.0005 | 1 | 0 | 0 | 40 | 0 | 5e-05 | 0.005 | -6.832 | 0.053 |  | 0.253 | -3.525 | 1.044 |  | 摸底

## case16 test report
dynamic_test | ADC_samp_freq | fund_wave_freq | ADC_OSR | scan_freq | PGA | fund_wave_amp_vol | fund_wave_amp_current | power_factor | vol_dc_bias | current_dc_bias | harmonic_wave_oder | amp_harmonic_vol | amp_harmonic_current | full_RMSU_err (RMSU电压有效值误差(‰)) | full_RMSI_err (RMSI电流有效值误差(‰)) | full_P_err (P有功功率误差(‰)) | full_Q_err (Q无功功率误差(‰)) | fund_RMSU_err (RMSU电压有效值误差(‰)) | fund_RMSI_err (RMSI电流有效值误差(‰)) | fund_P_err (P有功功率误差(‰)) | fund_Q_err (Q无功功率误差(‰)) | 备注
case16 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.1 | 1 | 0 | 0 | 5 | 0 | 0 | -0.018 | 0.12 | 0.021 |  | -0.058 | -0.052 | -0.125
case16 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 5 | 0 | 0 | -0.018 | 0.142 | 0.022 |  | -0.058 | -0.044 | -0.139
case16 | 819200 | 50 | 128 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 5 | 0 | 0 | -0.018 | 1.919 | 0.131 |  | -0.058 | -2.654 | -1.449 |  | 摸底
case16 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.1 | 1 | 0 | 0 | 5 | 0 | 0 | 0.034 | 0.394 | 0.25 |  | 0.253 | 0.244 | 0.525
case16 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 5 | 0 | 0 | 0.034 | 0.389 | 0.252 |  | 0.253 | 0.233 | 0.527
case16 | 819200 | 50 | 256 | 50 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 5 | 0 | 0 | 0.034 | 0.019 | 0.432 |  | 0.253 | -1.242 | 0.784 |  | 摸底
case16 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.1 | 1 | 0 | 0 | 5 | 0 | 0 | -0.018 | 0.12 | 0.021 |  | -0.059 | -0.053 | -0.127
case16 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 5 | 0 | 0 | -0.018 | 0.142 | 0.022 |  | -0.059 | -0.046 | -0.141
case16 | 983040 | 60 | 128 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 5 | 0 | 0 | -0.018 | 1.927 | 0.131 |  | -0.059 | -2.655 | -1.452 |  | 摸底
case16 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.1 | 1 | 0 | 0 | 5 | 0 | 0 | 0.034 | 0.394 | 0.25 |  | 0.253 | 0.244 | 0.524
case16 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.05 | 1 | 0 | 0 | 5 | 0 | 0 | 0.034 | 0.388 | 0.252 |  | 0.253 | 0.233 | 0.527
case16 | 983040 | 60 | 256 | 60 | 1 | 0.4 | 0.001 | 1 | 0 | 0 | 5 | 0 | 0 | 0.034 | 0.02 | 0.433 |  | 0.253 | -1.241 | 0.784 |  | 摸底